Table of Contents

# Change Control and Document Register

Document: ECO-EEG-016 Revision: C Date: 1 September 2026, corrected 2 September 2026 Issued by: TI One Voice research programme (one.witysk.org), Brussels, Belgium Licence: CC BY-SA 4.0 Governing documents: DSN-EEG-003 Rev D, then RFQ-EEG-001 Rev E. Where this document and tools/design.py disagree, design.py governs.

Revision note, Rev A to Rev B. The routing-scope change is renumbered from ECO-EEG-016 to ECO-EEG-018 so that no change shares a number with this document; nine further changes (ECO-EEG-019 to ECO-EEG-027) are recorded; the enlargement of the carrier to 150.0 × 130.0 mm and its move from two layers to four are recorded inside ECO-EEG-018, which is the change that found them; the document register is rewritten with the revision letter of every document in the 1 September 2026 correction release, with the document identifiers that the package cites but never registered, and with the statement, which Rev A omitted, that nothing has been built.

Corrections within Rev B, made on 1 September 2026, closing the findings of the second cross-document audit that name this document. RUL-EEG-021 and SIM-EEG-018 are registered in section 1.1; every revision letter this document states is brought to the letters of that section and the escape clause that let citations drift is withdrawn; the numbers ECO-EEG-028 and ECO-EEG-029 are withdrawn and their two findings are recorded inside ECO-EEG-018, which is where RUL-EEG-021 puts them; the claim that the board is DRC-clean is replaced by what kicad/EEG-CAR-01_RevB_DRC_report.txt actually reports; and the MP-01 slot geometry is corrected to the 12 × 3 mm that tools/mech_gen.py cuts. The revision letter is deliberately not advanced: these are corrections inside the same release.

Correction within Rev B, made on 2 September 2026: the board closed. The DRC now reports zero violations and all 145 nets connected, so the fabrication-release gate of section 3 is met and that section is rewritten to say so. ECO-EEG-018’s routed-result table carries the current figures and section 2B records what was changed to close the board and what closing it cost. That paragraph said the fabrication data was released for review under RFQ-EEG-002A and not for fabrication, because no human layout engineer had read the routing. It is superseded at Rev C: one has read it, declined the review and advised a redesign, and Rev B is withdrawn from fabrication – ECO-EEG-030. The revision letter was not advanced for the 2 September correction: it was a correction inside the same release.

Further corrections within Rev B, made on 2 September 2026, closing the findings of an independent external review of the package that name this document. TOOL-EEG-022 Rev A is registered in sections 1.1 and 1.5, having shipped, and been cited by two released documents, while this register said that a document not on its list is not part of the release; the three workbooks in docs/ and the built webtest page are named in section 1.2 as the generated artifacts they are; RFQ-EEG-002A and RFQ-EEG-002B are ruled in section 1.3 as scopes of work rather than documents. Every production-simulation total this document states is brought to the run of 2 September 2026 – 171 passed, 0 failed, 5 open at the time this note was written; superseded later the same day by 193, 0 and 6, see section 1.1 – and the total in the C-01 row of section 2A is marked as the history it is. ECO-EEG-019’s claim that the Sallen-Key corner tolerance in TST-EEG-004 “is widened to match” the X7R substitution is withdrawn: no such limit existed to widen, and E-11’s low-pass half was not met with the approved parts – it is met later the same day on C0G parts, which is recorded under ECO-EEG-019 and does not revive the withdrawn sentence. Section 2D is added, recording the firmware and test-tool defects found and closed on the same day. The revision letter is again not advanced: these are corrections inside the same release.

Further corrections within Rev B, made on 2 September 2026, after the firmware was built and after four requirements moved. Section 1.6 is added, registering the four released trees that were on no list at all – fixtures/, records/, kicad/wh-bus-01/ and the provisioning and ATECC tooling under firmware/tools/. That is the defect of section 1.3 reappearing in the register for the second time, now against shipped data rather than against a document, while section 1.1 went on saying that what is not on its list is not part of the release. Section 1.2’s claim that everything under kicad/ is produced by emit_all.py from design.py is corrected: kicad/wh-bus-01/ is written by tools/wh_bus.py, which imports nothing from design.py and is not called by emit_all.py. Every production-simulation total this document states is brought to the run of 2 September 2026 – 193 passed, 0 failed, 6 open – and the claim that the simulator wrote this register’s copy of that total is withdrawn: it writes SIM-EEG-018 and nothing here, which is exactly why the figure could drift and did. Section 2E is added, recording the first real ESP-IDF build, the five defects it found, the emulator run and the provisioning opcode collision. ECO-EEG-024 is implemented, so S-02 is met in the design at 36.8 µA; E-11’s low-pass half is met on C0G parts; E-27’s phase driver is written; and the bias-lead topology finding of section 2C is fixed. Each of those is corrected where this document states it, with its date, rather than deleted. The revision letter is again not advanced: these are corrections inside the same release.

Nothing in this package has been manufactured, assembled or measured, and no safety engineer has reviewed this design. Every figure marked calculated is a calculation. Every ECO below that is marked implemented is implemented in the design source, not in copper. That statement is repeated in every controlled document on purpose.

The firmware is the single exception to “nothing has been built”, and it is a narrow one. Since 2 September 2026 it compiles against a real ESP-IDF – v5.2.5, target esp32s3 – and the four images in firmware/release/ with the SHA-256 manifest beside them are that build; it has also run one boot cycle under QEMU. It has never run on hardware, because no unit exists, and the emulator carries none of the peripherals the firmware talks to. Built is not run, and run in an emulator is not brought up. Section 2E states what that build and that run do and do not prove.

## Why this document exists

Package v1 had no change control and no document register. Revisions were recorded inside each document’s own front matter, so nothing said which revisions belonged together, nothing recorded why a change was made, and a manufacturer holding four files had no way to tell whether they were the same release. Worse, the same part identifier meant two different parts in two documents and nobody noticed, because there was no register to notice it in.

Package v2 Rev A fixed the part-identifier namespace and then reproduced the same failure in its own: it numbered one of its changes ECO-EEG-016, which is also this document’s number, and it registered twenty-one documents while the package went on to cite eighteen more that were on no list at all, six of them colliding numerically with documents that are. Rev B closes both.

This document does four things: it registers every controlled document and its current revision; it names the identifiers the package cites that are not controlled documents, and says what they are instead; it records every engineering change from package v1 to v2, with the reason, the verification and the impact; and it states the procedure for the next change.

## 1. Document register

### 1.1 Controlled documents

Every controlled document in package v2, at the revision released on 1 September 2026. A document not on this list is not part of the release. Section 1.3 deals with the identifiers that other documents cite and that are not on this list, and section 1.6, added on 2 September 2026, deals with the released data sets that carry no document identifier at all. Until section 1.6 existed, four substantial trees – fixtures/, records/, kicad/wh-bus-01/ and the provisioning tooling under firmware/tools/ – shipped inside a release whose own register said that what is not listed is not part of it.

| ID | Title | Rev | Status | Governs |
| --- | --- | --- | --- | --- |
| DSN-EEG-003 | Manufacturing design package | D | released | architecture, carrier, precedence – the governing document |
| RFQ-EEG-001 | RFQ and engineering specification | E | released | requirements and acceptance (M/S/O), pricing |
| ICD-EEG-006 | Interface control document | C | released | module interfaces, jumper schedule, GPIO map |
| SCH-EEG-005 | Schematic set, 8 sheets | B | generated | the circuit |
| DSN-EEG-002 | Helmet design and assembly | E | released, not reissued in this round | helmet, electrodes, wiring, fitting, case |
| PARTS-EEG-019 | Part identifier register | B | released | every part number in the package |
| ASM-EEG-007 | Assembly work instructions | B | released | how one unit is built |
| WH-EEG-008 | Harness and cable assembly | B | released | every cable, every conductor |
| TST-EEG-004 | Production test specification | C | released | what every unit is tested to, and it owns the T-numbers |
| JIG-EEG-009 | Test fixture design | C | released | the fixtures TST-EEG-004 needs, FIX-01 to FIX-04 |
| QP-EEG-010 | Quality plan | B | released | IQC, FAI, AQL, records, traceability |
| RISK-EEG-011 | Risk analysis and safety review pack | B | released as an input | the pack the safety reviewer receives |
| REG-EEG-012 | Regulatory and compliance file | B | released | RoHS, REACH, UN38.3, CISPR, 60601 gap analysis |
| SVC-EEG-013 | Service and refurbishment manual | B | released | the turnaround between participants |
| IFU-EEG-014 | Participant quick-start and placement guide | B | released | what goes in the case lid |
| PKG-EEG-015 | Packing, labelling and shipping | B | released | packing list, labels, foam, lithium shipping |
| ECO-EEG-016 | Change control and document register | C | released | this document |
| AVL-EEG-017 | Approved vendor list | C | released | sourcing, alternates, substitution |
| SIM-EEG-018 | End-to-end production simulation report | A | released, generated | the dry run of the manufacturing route against the data package |
| FW-EEG-001 | Firmware build and provisioning | D | released | firmware, build, provisioning, host tool, key fingerprint |
| MECH-EEG-020 | Printed part drawings | A | generated | dimensioned drawings of every printed part |
| RUL-EEG-021 | Rulings register | B | released | one answer to each cross-document disagreement, cited by section letter |
| TOOL-EEG-022 | Connectivity test program: specification, technical description and user manual | A | released | the browser tool that tests the USB link, and the browser client the study reuses |
| LAY-EEG-034 | Carrier layout rule sheet | A | released, generated, issued 8 September 2026 | the rule set the external layout contractor works under, and part of the order specification. Written by tools/emit_rule_sheet.py from tools/rules.py, so it and kicad/EEG-CAR-01_RevC.kicad_dru cannot disagree |
| ASM-EEG-023 | Register of programme assumptions | B | released, issued 2 September 2026 | every open question the completeness audit left that was a DECISION rather than a missing file, each decided, attacked by an independent reviewer and re-decided against that attack. Twenty-three entries at Rev B, eleven of which block a build, three of which cannot be decided from what exists. The twenty-third, MECH-D6-MODULE-ENVELOPES, is new under ECO-EEG-033 and blocks a build. Nothing in it is a sign-off |

Four entries need reading carefully rather than counting.

SIM-EEG-018 is a report written by tools/simulate_production.py on each run. The run behind this issue prints 193 passed, 0 failed and 7 open, and that is what its own file states because the simulator wrote it. It is not the same 193 as the run of 2 September 2026. Six checks were added under ECO-EEG-030 to ECO-EEG-033 – the Rev C input set assembled, the schematic matching the design source, every footprint audited, Rev B regraded, the rule sheet issued, the collision check run – and six were lost, because cadquery could not be installed on the machine this round was run on, so the HM-04/HM-05B bayonet fit and the HM-01P channel section were not measured. The two happen to cancel. A passed count is a count of the checks that RAN, and a run on a machine with cadquery would print 199. It is a file in docs/, but it is never edited by hand: it is the output of the last run, and a correction to it is a correction to the simulator or to design.py. The last run, of 8 September 2026, reports 193 checks passed, 0 failed and 7 open items, and docs/SIM-EEG-018_RevA_production_simulation.md states those three figures because the simulator wrote that file. It did not write this one. Rev B said the totals agreed “because the simulator wrote both”; that is false and is withdrawn on 2 September 2026. The copy in this section is typed by hand from the run, which is precisely how it came to stand at 171 while the report had moved to 193, and it is checked against the report rather than trusted.

The six open items of that run are: the static IRAM pool, which links with one byte free of 16,384; the carried-over v1 HM-01 mesh, which is watertight and is two disconnected bodies; that no human layout engineer has read the routing, which is the fabrication-release gate of section 3; that E-27 has never been seen to light – the phase driver is written and the current budget is met, but no unit exists, so TST-EEG-004 T11 has not been run; the unreconciled board current; and that SR-01 is closed in the design and not yet signed off – S-02 is met at 36.8 µA because ECO-EEG-024 was implemented, and applying the fix an analysis pointed to is not the same as having it approved.

Four figures, four dates, and only the last one is the total. 169 passed, 0 failed, 5 open at the 1 September 2026 issue, whose first open item was the 25 DRC violations; 171, 0 and 5 earlier on 2 September; 193, 0 and 6 at the last run of 2 September, which is what Rev B stated and is now history; 193, 0 and 7 at the run behind this release, on 8 September 2026. The seventh open item is that there is no fabrication data for the current board at all: Rev B is withdrawn and Rev C is unrouted. The one that was there and has changed its words is the layout review – a human layout engineer HAS now read the Rev B routing, which is why Rev B is withdrawn, and what is open is the review of the CONTRACTOR’S PLACEMENT of Rev C, which cannot happen until there is a placement.

The two 193s are not the same 193 and the coincidence is worth a sentence, because section 1.1’s rule is that a passed count moves when the check set moves. Six checks were added in this round and six were lost, and they cancel: added are the Rev C input set assembled, the schematic matching the design source, every footprint audited against its part number, Rev B regraded under the ECO-EEG-032 rule set, the rule sheet issued and the collision check run; lost are the six that need cadquery, which could not be installed on the machine this round ran on, so the HM-04/HM-05B bayonet fit and the HM-01P channel section were not measured. A run on a machine with cadquery would print 199. Between the two runs of 2 September, one item left the list – E-11’s low-pass half, because the Sallen-Key moved to C0G; two joined it – the IRAM pool and the HM-01 mesh, both found by measuring things nobody had measured; and two were restated rather than closed – E-27’s moved from the phase driver is not written to it has never been seen to light, and S-02’s from the single-fault current fails to SR-01 is closed in the design and not signed off. The DRC violations had left the list before either run, when the report reached zero. The passed count moves because the check set moves and not because anything is graded more leniently. The failed count is 0 in every run. A total quoted anywhere in this package that is not the total of the last run is stale, and the total to quote is whatever tools/simulate_production.py last printed, not a figure copied between documents – including this paragraph, which is a copy like any other.

RUL-EEG-021 began as an internal audit worksheet, tools/RULINGS.md. Four released documents came to cite it, which made an uncontrolled file part of the release by the back door. It is now a controlled document at Rev A and is cited as RUL-EEG-021 with a section letter – section A for the geometry taken from the source files, section B for the instrument, section C for the structural rulings. The worksheet under tools/ is not the controlled copy and is not cited by any released document.

DSN-EEG-002 Rev E is the helmet document. It was not reissued in this correction round, so it still carries the package v1 statement that the microphone preamplifier is on the boom, which RUL-EEG-021 section B and ICD-EEG-006 section 2.9 correct to the MP-01 module plate at J21. Until the next issue of DSN-EEG-002, ICD-EEG-006 governs on that point and DSN-EEG-002 is wrong on it.

TOOL-EEG-022 is the connectivity test program: one self-contained HTML page, webtest/EEG-Connectivity-Test.html, built by webtest/build.py from the modules under webtest/js/. It is the browser end of the wire format, and the study’s own client is the same protocol.js, transport.js and sinks.js with a different sink, so it is production code and not a throwaway page. It was added to this register on 2 September 2026, having been written, issued at Rev A and cited by other documents while appearing on no list at all. That is the defect ECO-EEG-015 was raised against in the part namespace and section 1.3 was written to close in this one, reappearing in the register itself: this section states that a document not on this list is not part of the release, so for one release TOOL-EEG-022 was formally not part of a release it shipped in. FW-EEG-001 Rev D section 5 is normative for the wire format and governs where the tool and the specification disagree, and TST-EEG-004 owns the T-numbers under which the tool is run – T30, the host link check – so a step that uses the tool is cited from TST-EEG-004 and not from TOOL-EEG-022.

### 1.2 Generated artifacts

Everything under kicad/, kicad/gerber/, schematic/ and graphics/ is produced by tools/emit_all.py from tools/design.py – with one exception, corrected here on 2 September 2026: kicad/wh-bus-01/. That directory is the WH-BUS-01 Rev A fabrication set for the contact-light bus board, and it is written by tools/wh_bus.py, which emit_all.py does not call and which imports nothing from design.py, gerber.py, pcbgen.py or router.py. That independence is deliberate – WH-BUS-01 is ten pads and one copper bar, and running it through a 211-part four-layer board’s router and DRC would couple two boards that have no reason to be coupled – but it means re-running emit_all.py does not regenerate WH-BUS-01, and does not prove kicad/ current. Everything else in the rule stands: these files are not edited by hand, and if one of them is wrong the source is wrong and the fix goes in the source.

The mechanical trees follow the same rule under tools that emit_all.py does not call: mech/stl and mech/step from tools/mech_gen.py, with the board-outline STEP and the carried-over HM-01 STL from tools/emit_extras.py; mech/drawings from tools/mech_drawings.py; mech/renders from tools/assembly_render.py. A full reissue is therefore several commands, not one, and emit_all.py alone does not prove the mechanical set current. firmware/main/board_pins.h is declared generated from design.py and is not: no tool under tools/ writes it, and tools/simulate_production.py only reads it back and checks it against design.py. It is maintained by hand against that check, and that is an open item.

Three more trees are generated by tools emit_all.py does not call, and until 2 September 2026 no section of this register named any of them: fixtures/ from tools/fixture_gen.py (except fixtures/firmware/, which is hand-written), records/ from records/make_records.py, and kicad/wh-bus-01/ from tools/wh_bus.py. Section 1.6 registers all three with the document that owns each.

The three workbooks in docs/ are in this class and were named nowhere in this register until 2 September 2026. EEG_kit_BOM_for_bidders_RevC.xlsx, EEG_kit_BOM_INTERNAL_RevC_costed.xlsx and EEG_kit_manufacturer_contacts_RevB.xlsx are written by tools/emit_workbooks.py, which calls tools/emit_costed_bom.py for the second of them, and emit_all.py does not call any of it. They carry revision letters in their file names, so they read as controlled documents and are not: they are output, they are not edited by hand, and section 1.4 ranks the kit BOM workbook last in precedence for that reason. The same applies to webtest/EEG-Connectivity-Test.html, which webtest/build.py inlines from the modules under webtest/js/ and which TOOL-EEG-022 owns.

The Rev C layout-input set is in this class and is registered here on 8 September 2026 (ECO-EEG-033). tools/emit_all.py no longer regenerates the Rev B set – it refuses, and says why (ECO-EEG-031) – and writes these instead, all from design.py and none edited by hand:

| File | Written by | What it is |
| --- | --- | --- |
| kicad/EEG-CAR-01_RevC.kicad_sch and its nine sheets | tools/emit_kicad_sch.py | the native hierarchical schematic, KiCad 8 |
| kicad/EEG-CAR-01.kicad_sym | tools/emit_kicad_sch.py | the symbol library those sheets are drawn from |
| kicad/EEG-CAR-01_RevC_schematic_netlist_check.txt | tools/sch_netlist.py | the schematic read back and diffed against design.py |
| kicad/EEG-CAR-01_RevC.kicad_dru | tools/rules.py | the custom design rules, 37 of them |
| kicad/EEG-CAR-01_RevC_BOM.csv | tools/gerber.py | the BOM, with the substitution column |
| kicad/EEG-CAR-01_RevC_CPL_SMT_top.csv, ..._CPL_THT_top.csv | tools/gerber.py | provisional placement; the files say so |
| kicad/EEG-CAR-01_RevC-IPC-D-356A.ipc | tools/gerber.py | the netlist, from the pads alone |
| kicad/EEG-CAR-01_RevB_regraded_ECO-EEG-032.txt | tools/grade_revb.py | the released Rev B geometry under the new rule set |
| docs/footprint_audit_RevC.md | tools/footprint_audit.py | every designator’s footprint against its part number |

SCH-EEG-005 is in this class. Its revision letter follows the board revision letter in design.py, which is Rev B, and it is regenerated rather than edited whenever the netlist changes. It is a picture and not a netlist, which is finding 7 of ECO-EEG-030; the netlist-bearing schematic is the Rev C set above, and the two are emitted separately for now because tools/schematic.py still draws the Rev B PDF. The board is still Rev B after ECO-EEG-019 to ECO-EEG-027, because Rev A was the package v1 carrier and Rev B is the package v2 carrier throughout its drafting.

### 1.3 Identifiers the package cites that are not controlled documents

Package v2 Rev A registered twenty-one documents and stated that nothing else is part of the release, and then the seventeen documents went on to cite eighteen further identifiers as though they were controlled. Six of those collide numerically with registered documents, which is exactly the failure ECO-EEG-015 was raised to fix in the part namespace. This register owns the document namespace, so it rules on them here.

Withdrawn, because they collide with a registered document. The content is a numbered section of the document that already owns it, and is cited that way.

| Withdrawn identifier | Collides with | Cite instead |
| --- | --- | --- |
| REL-EEG-009 | JIG-EEG-009 | FW-EEG-001 section 9 (release and versioning) |
| PROV-EEG-007 | ASM-EEG-007 | FW-EEG-001 section 7 (provisioning) |
| KEY-EEG-008 | WH-EEG-008 | FW-EEG-001 section 7 (key handling and fingerprint) |
| SD-EEG-013 | SVC-EEG-013 | FW-EEG-001 section 10 (microSD file format) |
| OTA-EEG-014 | IFU-EEG-014 | FW-EEG-001 section 9 (field update) |
| CAL-EEG-012 | REG-EEG-012 | QP-EEG-010 section 12 (calibration record) |
| ETH-EEG-001, DIS-EEG-001, KPL-EEG-001, QSC-EEG-001, PLG-EEG-001 | RFQ-EEG-001 | the research documents of section 4, PKG-EEG-015 section 1 (packing list) and PKG-EEG-015 section 4 (quick-start card) as applicable |

Not documents at all. ART-LBL-01 to ART-LBL-07, ART-PACK-01, ART-DIS-01, ART-RET-01 and DRW-LBL-PLACEMENT are label and packaging artwork files controlled as generated artifacts under graphics/ by PKG-EEG-015 section 5. QF-EEG-006-01 is a quality form, controlled by QP-EEG-010 section 8. They keep their identifiers as file names; they are not registered as documents and they carry no revision letter of their own, only the revision of the document that owns them.

RFQ-EEG-002A and RFQ-EEG-002B are scopes of work, not documents. RFQ-EEG-002A is cited in fifteen of the documents in docs/, and it is what the fabrication-release gate of section 3 waits on; until 2 September 2026 no register named it at all. They are defined in RFQ-EEG-001 Rev E section 1.1: 002A is the review and sign-off of the supplied carrier routing, which is the gate of section 3 below, and 002B is the prototype builds and complete field kits in three phases. Neither has a file, a revision letter or an author, and neither may be cited as though it had one; the document that carries them is RFQ-EEG-001 at the letter in section 1.1, and a statement about what 002A covers is a citation of RFQ-EEG-001 Rev E section 1.1.

MECH-EEG-020 was cited before it was registered. It is now registered in section 1.1 as a generated artifact. TOOL-EEG-022 was cited before it was registered too: two released documents, FW-EEG-001 section 5 and RFQ-EEG-001 Rev E, named it while this register did not. It is registered in section 1.1 and section 1.5 as of 2 September 2026, and what it is is set out under section 1.1.

tools/RULINGS.md and the simulation report were cited before they were registered too, which is the same defect in a worse place: four released documents – TST-EEG-004 section 3, ASM-EEG-007 section 2.2, AVL-EEG-017 and this document at ECO-EEG-020 – cited an uncontrolled worksheet under tools/ as though it were a released document, and several cited the production simulation by name. Both are now controlled documents in section 1.1: the worksheet is issued as RUL-EEG-021 Rev B and is cited as RUL-EEG-021 with a section letter, and the simulation report is SIM-EEG-018 Rev A. The file tools/RULINGS.md is not a document, is not the controlled copy, and is not cited by anything in the release.

The unit serial-number format is not a document but it was claimed by three of them. It is ruled: TIOV-B-nnnn – programme prefix, hardware revision letter, four digits, Phase 1 using 0001 to 0009. It is defined once in PKG-EEG-015 section 5 and appears identically in the label text, the Data Matrix, the USB iSerialNumber, the calibration record and the packing list. Every other form that appeared in package v1 and in package v2 Rev A drafts is withdrawn.

### 1.3A One number pool for documents and for changes

Added 8 September 2026 (ECO-EEG-030). Document identifiers and change identifiers have always been drawn from the same EEG-nnn run, and the register has already been bitten once by that: ECO-EEG-016 is this document and was also, in a package v2 Rev A draft, a change number, so a citation of it was ambiguous everywhere it appeared. That change is now ECO-EEG-018 and 016 is reserved to the document. The same trap is open at every other number, because ECO-EEG-024 is a change and nothing stops a future document being numbered 024.

The rule from here: a number is allocated once, from one pool, whatever prefix uses it. The next free number is one above the highest number used by either namespace, counting withdrawn numbers as used. Numbers 024 to 029 are not available to a document: 024 to 027 are implemented changes and 028 and 029 are withdrawn ones.

Allocated under this rule on 8 September 2026: ECO-EEG-030, ECO-EEG-031, ECO-EEG-032 and ECO-EEG-033 as changes, and LAY-EEG-034 as a document – the carrier layout rule sheet issued under ECO-EEG-033. The next free number after them is 035.

### 1.4 Precedence

Highest first: DSN-EEG-003 → RFQ-EEG-001 → ICD-EEG-006 → SCH-EEG-005 → DSN-EEG-002 → PARTS-EEG-019 → the kit BOM workbook. Where a number appears in a document and in tools/design.py, design.py governs.

### 1.5 Revision letters in this release

Eight documents advanced a letter on 8 September 2026 under ECO-EEG-030 to ECO-EEG-033, and the set is at the letters in section 1.1: DSN-EEG-003 to Rev D; ECO-EEG-016, ICD-EEG-006, JIG-EEG-009 and AVL-EEG-017 to Rev C; FW-EEG-001 to Rev D; ASM-EEG-023 and RUL-EEG-021 to Rev B. LAY-EEG-034 is new at Rev A. The bump, the 218 citations that had to follow it and the file renames were done by tools/revision_bump.py rather than by hand, because doing it by hand across twenty documents is how this package came to have sixty disagreements in the first place. Three documents the round did not touch keep their letters, deliberately: PARTS-EEG-019 stays Rev B because ECO-EEG-031 changed no part identifier, only manufacturer part numbers, which AVL-EEG-017 owns; RISK-EEG-011 stays Rev B because none of the four ECOs touches the electrode path, the isolation, the battery, the charge interlock or the patient-current budget; and RFQ-EEG-001 stays Rev E, with E-26’s self-contradicting switch class recorded as an open item rather than corrected here.

Every document marked in bold in section 1.1 was reissued on 1 September 2026 and its revision letter advanced by one. RUL-EEG-021, SIM-EEG-018 and TOOL-EEG-022 are new, at Rev A; TOOL-EEG-022 was issued at Rev A on that date but was not registered here until 2 September 2026, which is a defect in this register and not a second issue of the tool. ASM-EEG-023 is also new at Rev A and was issued on 2 September 2026, one day after the rest, which is why it is the one document in section 1.1 whose date is not 1 September. The file names in docs/ carry the letters, and the letters in section 1.1 are the release.

A cross-reference names a document by its identifier and section, not by its revision. Where a cross-reference does give a letter, that letter must be the one in section 1.1:

| Document | Letter to cite |
| --- | --- |
| DSN-EEG-003, the governing document | Rev D |
| RFQ-EEG-001 | Rev E |
| TST-EEG-004 | Rev C |
| FW-EEG-001 | Rev D |
| DSN-EEG-002 | Rev E |
| ICD-EEG-006, JIG-EEG-009, ECO-EEG-016, AVL-EEG-017 | Rev C |
| SCH-EEG-005, PARTS-EEG-019, ASM-EEG-007, WH-EEG-008, QP-EEG-010, RISK-EEG-011, REG-EEG-012, SVC-EEG-013, IFU-EEG-014, PKG-EEG-015, ASM-EEG-023, RUL-EEG-021 | Rev B |
| SIM-EEG-018, MECH-EEG-020, TOOL-EEG-022, LAY-EEG-034 | Rev A |

An older letter in a cross-reference is a defect to be corrected, not a synonym. Rev A of this document said that an older letter “means the same document as corrected in this release”; that escape clause is withdrawn, because a package whose citations are allowed to drift cannot be checked against itself, which is the failure the whole register exists to prevent. A release is the whole set at the letters in section 1.1 or it is nothing.

Two letters are easy to trip over, and both are correct as written. The board is EEG-CAR-01 Rev B and stays Rev B through this correction round, so “an ECO against Rev C of the board” in section 2.3 means the next board revision and not DSN-EEG-003 Rev D. SCH-EEG-005’s letter follows the board’s letter, not the document round’s, which is why it is B while the documents around it are B or C for a different reason.

### 1.6 Released data sets that carry no document identifier

Added 2 September 2026. Four trees ship in this release, are relied on by released documents, and were named in no register at all. That is the failure of section 1.3 reappearing for the second time inside the register itself, and on a larger scale than the TOOL-EEG-022 case, because none of these is a single file: fixtures/ alone carries two fixture board sets, seven printed parts and a firmware project. They are not promoted to controlled documents – none of them is a document – they are registered here as released data, each with the document that owns it, the tool that writes it, and what it is not.

| Tree | What it is | Owned by | Written by | Status |
| --- | --- | --- | --- | --- |
| fixtures/ | the FIX-01 to FIX-04 test-fixture data: board data for pcb/FIX-01 and pcb/FIX-04, seven printed parts in step/ and stl/, the M1-M3 fixture controller firmware under fixtures/firmware/, and MANIFEST.json with a SHA-256 for every file | JIG-EEG-009 Rev C | tools/fixture_gen.py, except fixtures/firmware/, which is hand-written C | released as fixture data. The two fixture boards are NOT a fabrication set – there is no copper layer on either, and fixtures/README_fixture_data_index.txt says why |
| records/ | the machine-readable per-unit test record: the JSON schema, a worked example record, the lot-summary CSV header and the calibration-certificate template | TST-EEG-004 Rev C sections 12 and 13 | records/make_records.py, which reads the step list out of TST-EEG-004 at generation time and fails rather than writes if the document and the table disagree | released, generated |
| kicad/wh-bus-01/ | the WH-BUS-01 Rev A fabrication data for the contact-light bus board: Gerbers, drill, IPC-D-356A netlist, placement and BOM note, and the checksum file | PARTS-EEG-019 Rev B registers the part; WH-EEG-008 Rev B owns the harness it serves | tools/wh_bus.py – not emit_all.py; see section 1.2 | released, generated |
| mech/step/footprints/ and the three Rev C body files in mech/step/ | the 3D bodies of ECO-EEG-033: one per footprint class, bound into kicad/EEG-CAR-01_RevC.kicad_pcb; the carrier’s component height envelopes; MP-01 and its standoffs; and the thirteen module assemblies as declared maximum envelopes, which are assumptions (ASM-EEG-023 MECH-D6-MODULE-ENVELOPES) and not measurements | ICD-EEG-006 Rev C section 4 owns MP-01; ASM-EEG-023 owns the envelopes | tools/mech_bodies.py on tools/step_write.py – not tools/mech_gen.py, which needs cadquery and could not be run here | released as body data. NOT a printable part set: they are envelopes, not parts, and no STEP reader has read them |
| mech/EEG-CAR-01_RevC_collision_check.txt | the interference report of ECO-EEG-033 finding 6 | ASM-EEG-023 MECH-D6-MODULE-ENVELOPES | tools/collision_check.py | released, generated. Three of its four cases are open and one blocks a build |
| firmware/tools/ | the provisioning station: provision.py and provision_selftest.py, atecc608b_config.py and the three configuration files it writes, calibration_schema.py and its schema, and verify_stream.py | FW-EEG-001 Rev D section 7 | hand-written, except the three ATECC configuration files and the calibration schema, which their own generators write and check | released as tooling. ATECC608B_CONFIG_TEMPLATE.md is a PROPOSAL in its own words – not reviewed, never written to a part, not released for production |

Two files in that last tree read like documents and are not. firmware/tools/README_provisioning.md and firmware/tools/ATECC608B_CONFIG_TEMPLATE.md both open with a Document: line, a licence and a list of governing documents, which is the front matter of a controlled document in this package. They carry no revision letter and they are not controlled documents: they are the tooling’s own documentation, owned by FW-EEG-001 section 7, and where either disagrees with FW-EEG-001 the specification governs. Front matter does not make a document; section 1.1 does.

Where the boundary is. A tree in this section is part of the release and may be cited by a released document, by path and by the owning document’s section. It is not cited as though it had a revision letter, and a statement about what it contains is a citation of its owner in section 1.1. Anything shipped that is on neither list is the defect this section exists to stop, and the next one is raised against this section rather than discovered by an outside reviewer, which is how these four were found.

## 2. Engineering changes, package v1 to v2

Twenty-six changes: twenty-one to the carrier design, three to the package structure, and two that change a requirement without changing the carrier. That split is the register in section 2.1 counted: carrier is ECO-EEG-001 to ECO-EEG-014, ECO-EEG-019 to ECO-EEG-024 and ECO-EEG-027; package structure is ECO-EEG-015, ECO-EEG-017 and ECO-EEG-018; requirement-only is ECO-EEG-025 and ECO-EEG-026.

Eight of the twenty-six are classed blocking – ECO-EEG-001 to ECO-EEG-004, ECO-EEG-007 to ECO-EEG-009 and ECO-EEG-013 – and every one of them is a package v1 defect that meant the board could not have been built. ECO-EEG-018 is classed scope and carries two further blocking findings of its own, that the board as drafted at 130 × 124 mm on two layers could not have been routed; that is why it is the longest entry in the register. ECO-EEG-024 is classed blocking for Phase 2. Rev B recorded two changes as not implemented, ECO-EEG-023 and ECO-EEG-024; corrected on 2 September 2026, ECO-EEG-024 is implemented and only ECO-EEG-023 remains open. Both are still addressed to the safety reviewer, because taking a change in the design source is not the same as having it approved: ECO-EEG-024’s own row in RISK-EEG-011, SR-01, is closed in the design and unsigned.

ECO-EEG-016 is this document’s identifier and is deliberately not used as a change number. The change numbers run ECO-EEG-001 to ECO-EEG-015 and ECO-EEG-017 to ECO-EEG-027, which is exactly what the table in section 2.1 contains, what RUL-EEG-021 section B rules, and what DSN-EEG-003 section 9 and RFQ-EEG-001 read the register as containing. In package v2 Rev A the routing-scope change was numbered ECO-EEG-016, so a citation of “ECO-EEG-016” was ambiguous everywhere it appeared; it is now ECO-EEG-018.

ECO-EEG-028 and ECO-EEG-029 are withdrawn. A Rev B draft of this document briefly gave the board size and the layer count their own numbers. No ruling creates them, and RUL-EEG-021 section A treats both as findings of the routing work rather than as changes in their own right, so they are recorded inside ECO-EEG-018 under their own headings. Neither number is reused: as with the part identifiers of ECO-EEG-015, a withdrawn number stays withdrawn.

This register is the one home for the change list. DESIGN_FACTS section 7 and DSN-EEG-003 section 9 index the same fourteen changes in their own phrasings; both are indexes, this section is the register, and where they differ this section governs.

### 2.1 The register

| ECO | Class | Subject | Status |
| --- | --- | --- | --- |
| ECO-EEG-001 | blocking | contact lights had no driver | implemented |
| ECO-EEG-002 | blocking | nothing supplied the 3.3 V rail | implemented |
| ECO-EEG-003 | blocking | there was no charge input | implemented |
| ECO-EEG-004 | blocking | the envelope filter was not in circuit | implemented |
| ECO-EEG-005 | major | rectifier diode pair the wrong way round | implemented |
| ECO-EEG-006 | major | filter in the wrong place in frequency | implemented |
| ECO-EEG-007 | blocking | the board had no mounting holes | implemented |
| ECO-EEG-008 | blocking | the DevKit would not have fitted | implemented |
| ECO-EEG-009 | blocking | the pin map used the PSRAM pins | implemented |
| ECO-EEG-010 | major | four functions declared and not connected | implemented |
| ECO-EEG-011 | major | no analogue reference at the module connectors | implemented |
| ECO-EEG-012 | major | non-plated holes would have been plated | implemented |
| ECO-EEG-013 | blocking | analogue module connectors could not be escaped | implemented |
| ECO-EEG-014 | major | light lines ran through the electrode harness | implemented |
| ECO-EEG-015 | major, documentation | part identifiers meant two different things | implemented |
| ECO-EEG-017 | major, compliance | lithium shipping was never mentioned | implemented |
| ECO-EEG-018 | scope, with two blocking findings | the routing is now supplied, and doing it grew the carrier to 150.0 × 130.0 mm and took it from two layers to four | implemented |
| ECO-EEG-019 | major | C1–C16 carried an X7R part number under a C0G requirement | implemented |
| ECO-EEG-020 | minor | the board had no fiducials | implemented |
| ECO-EEG-021 | major | the I²C bus had no pull-ups on the carrier | implemented |
| ECO-EEG-022 | major | VBUS_DET did not reach a guaranteed logic high | implemented |
| ECO-EEG-023 | major | ENV_CMP swings ±2.5 V into a 3.3 V pin | OPEN, not implemented |
| ECO-EEG-024 | blocking for Phase 2 | S-02 single-fault DC fails at 53.2 µA | implemented 2 September 2026, was OPEN; S-02 now met in the design at 36.8 µA and not signed off |
| ECO-EEG-025 | major | F-06’s three-minute ring buffer does not fit the PSRAM | implemented |
| ECO-EEG-026 | major | E-04’s −100 dB crosstalk is neither achievable nor measurable | implemented |
| ECO-EEG-027 | major | the envelope AC-coupling corner removed the envelope | implemented |
| ECO-EEG-028 | – | withdrawn; the board size is a finding of ECO-EEG-018 | withdrawn |
| ECO-EEG-029 | – | withdrawn; the layer count is a finding of ECO-EEG-018 | withdrawn |
| ECO-EEG-030 | scope, with seven findings | an external layout engineer read the Rev B board, declined the review and advised redesign; Rev B is withdrawn from fabrication and Rev C is opened as a reviewable design whose layout is bought | implemented 8 September 2026 |
| ECO-EEG-031 | major | four part numbers could not be bought or could not be fitted: the quad OPA4376 has no SOIC-14, the ferrite bead is not a Murata part, the tactile switch is a 12 x 12 mm part on a 6 x 6 mm land, and the patient-connected DIN sockets were not marked non-substitutable | implemented 8 September 2026 |
| ECO-EEG-032 | major | the DRC reported zero violations against a rule set that did not contain six of the seven findings; the rules are now encoded once and exported to KiCad, and Rev B is regraded under them | implemented 8 September 2026 |
| ECO-EEG-033 | scope | the Rev C reviewable design set: a native KiCad schematic, an unrouted board and project, 3D bodies and a collision check, and the layout rule sheet LAY-EEG-034 | implemented 8 September 2026 |

### ECO-EEG-001 – the contact lights had no driver

Class: blocking. Found by: netlist connectivity check (single-pad nets). Was: the eight LED1–LED8 nets existed on the harness connector J14 and nowhere else. The 74HC595 module socket J19 carried only power, ground, the three shift-register control lines and a nominal LED supply pair; the module’s eight Q outputs were not brought to the carrier at all. The contact lights – which the whole self-fitting design depends on, and which DSN-EEG-002 section 5 argues for at length – could not have worked. Now: J19 is a 1×16 socket exposing Q0–Q7 as well as OE and MR. Q0–Q7 drive the harness lines through R70–R77 (1 kΩ), giving (3.3 − 2.0) / 1000 = 1.3 mA per site and 10.4 mA total from GPIO48. The three colours come from eight lines and one common by driving two-lead bicolour LEDs in two phases at a nominal 240 Hz (see FW-EEG-001 and SCH-EEG-005 sheet 8). Corrected 2 September 2026: the driver that now exists alternates on the FreeRTOS tick, so the real rate is about 250 Hz rather than exactly 240; both halves are equal, so the duty is 50/50 and the colour does not shift. E-27 is written against “above 100 Hz” and 250 Hz meets it. The rate is stated rather than rounded back to 240. Verified: every LEDn net is one connected copper island in the DRC connectivity check; the drive current is computed in tools/simulate_production.py station 9. Impact: J19 pin count, eight new resistors, a firmware change, and a change to the bicolour LED part in the kit BOM. Was not closed by this ECO, and was closed on 2 September 2026: Rev B recorded that the bicolour phase scheme was specified and the driver was not written – lights_write() and lights_task() were on/off only, so E-27’s amber state had no implementation and TST-EEG-004 T11 could not pass. The driver is now written. Both halves of the ADS1299 lead-off word are read, where only LOFF_STATP was before: one bit per site cannot express three colours, which is where amber had been lost. Neither detector gone is green, exactly one is amber, both is red; green in phase A, red in phase B, amber in both. What is still open is different and smaller, and it is open item 4 of the simulation run: E-27 has never been seen to light. No unit exists, so nothing has been driven and T11 – which reads the red-to-green ratio with a colorimeter – has not been run. A driver that compiles is not a light that is amber. FW-EEG-001 owns the FW-Dnn status.

### ECO-EEG-002 – nothing supplied the 3.3 V rail

Class: blocking. Found by: netlist connectivity check. Was: DVDD3V3 had twelve pads, all of them module supply inputs, and no source. The charger module’s SYS output, the ESP32-S3-DevKitC-1’s 5 V pin and both ADS1299 module supply pins were unconnected nets. The board could not have powered up. Now: the charger’s power-path output is VSYS; a TPS63020-class buck-boost module at J25 makes V5V; V5V feeds the DevKit at J6.21 and both ADS modules at J1.11 and J3.11; the DevKit’s on-board regulator makes DVDD3V3. R86 pulls the buck-boost enable to VSYS so the rail starts before DVDD3V3 exists. C70–C74 are the bulk decoupling. Verified: power tree drawn on SCH-EEG-005 sheet 7; every rail is one connected island. Impact: one new module in the BOM, one new socket, six new passives, a change to the pricing template. Not closed by this ECO: the carrier draws a calculated 288 mA worst case from the DevKit’s on-board regulator, which is inside its rating but dissipates about 0.5 W inside a closed pod, so the rail is not proven. Phase 1 measures it at TST-EEG-004 T3 and reports the case temperature; see section 2.3.

### ECO-EEG-003 – there was no charge input

Class: blocking. Found by: netlist connectivity check. Was: VBUS_CHG had one pad, on the charger module socket. Nothing brought power to it. RFQ-EEG-001 Rev C E-23 called for “a USB-C connector for charging and data”, which cannot be reconciled with S-03’s requirement that no conductive path exist from the host connector to the patient side. Now: two separate connectors. Data is the USB-C on the ADuM4160 module and carries no power to the instrument. Charging is a separate charge-only receptacle wired to J24 through F1 (1.1 A PTC) and D24 (transient suppressor), with a divider (R84/R85) giving VBUS_DET on GPIO46. RFQ E-23 has been rewritten to say so and to carry the 45 °C charge inhibit – “charger IC with thermal regulation; no charging above 45 °C” – which is E-23, not S-04. Verified: two independent interlocks are described and tested in TST-EEG-004; the current path is drawn on SCH-EEG-005 sheet 7. Impact: one panel part, one pigtail, an enclosure opening, a label warning, and a line in the participant card. Not closed by this ECO: the divider value was wrong and is corrected by ECO-EEG-022. The thermistor of S-04 is not fitted and S-04 is not met – there is no NTC net in design.py and no thermistor way on J12 or J13 – so the 45 °C inhibit rests on the charger IC’s own thermal regulation alone. DSN-EEG-003 section 11 and RISK-EEG-011 carry it as an open hardware item. Also not closed: the host connector on the named ADuM4160 module is USB-B while E-24 asks for USB-C, which is a live non-conformance. The interim answer is a short USB-B-to-USB-C panel pigtail, WH-09, until an isolator module with a USB-C host connector is qualified. The host connection is a socket, not a captive cable; the captive lead through a gland is a Phase 2 item.

### ECO-EEG-004 – the envelope filter was not in circuit

Class: blocking. Found by: netlist connectivity check (single-pad nets ENVn_F_INM). Was: each envelope channel used one dual OPA2376. The second amplifier’s inverting input was unconnected and its non-inverting input was tied to AGND_REF, so the Sallen-Key network around it drove nothing and the amplifier’s output was an unconstrained buffer of the reference. The 50 Hz filter that RFQ E-11 requires did not exist electrically. Now: one OPA4376 quad per channel: A is the precision half-wave rectifier, B is the inverting summer that makes the absolute value, C is the Sallen-Key low-pass with both loops closed, D is the output buffer after the divider. Verified: SCH-EEG-005 sheet 3 draws every connection; simulate_production.py computes the corner frequency and Q from the fitted values. Impact: U1–U3 change from SOIC-8 to TSSOP-14 and from dual to quad; three new resistors per channel.

### ECO-EEG-005 – the rectifier diode pair was the wrong way round

Class: major. Found by: review of the netlist against the BAT54S datasheet pinout. Was: D20.1 was the op-amp output, D20.2 the summing node and D20.3 the rectified node. On a BAT54S, pin 3 is the common point of the series pair, so this connects two anodes into one node rather than forming a feedback rectifier. Now: pin 3 is the op-amp output; pin 1 is the rectified node and pin 2 the inverting input. The loop closes for both polarities, which is what makes it a precision rectifier. Verified: SCH-EEG-005 sheet 3; the topology is stated in the drawing note. Impact: netlist only. Same part, same footprint. Scope note: BAT54S is used only at D20, D40 and D60, the envelope rectifiers. D1–D16, the electrode clamps, are BAV99: Schottky leakage across the series resistor – 47 kΩ when this was written, 68 kΩ since ECO-EEG-024 was implemented on 2 September 2026, which makes the same argument harder, not easier – is an offset error on a 10 µV input. TST-EEG-004 T1 is corrected to say so.

### ECO-EEG-006 – the filter was in the wrong place in frequency

Class: major. Found by: computing the transfer function from the fitted values. Was: R = 33 kΩ, both capacitors 100 nF. Equal R and equal C in a unity-gain Sallen-Key gives Q = 0.5 and a −3 dB point at 0.64 × f₀ = 31 Hz, not the 50 Hz ± 10 % of RFQ E-11. Now: R = 22 kΩ, C to ground 100 nF, feedback C 220 nF → f₀ = 48.77 Hz, Q = 0.7416 (Butterworth), −3 dB at ≈ 50 Hz. The group delay at DC is 1 / (Q × 2π × f₀) = 4.40 ms, which is 4.40 samples at 1000 Hz; the √2 / (2π f₀) form that gives 4.61 ms is only correct at Q = 0.7071 and is withdrawn. Verified: computed in simulate_production.py station 9 and checked against E-11; JIG-EEG-009 section 2.3 carries the derivation. Impact: three resistor values and three capacitor values.

### ECO-EEG-007 – the board had no mounting holes

Class: blocking. Found by: comparing the specification sheet with the board file. Was: EEG-CAR-01_PCB_spec_sheet.txt promised “4 × M3 mounting holes (see placement PDF)” and DSN-EEG-003 Rev A.2 section 3 said “400 of 406 pads; the six unassigned are mounting holes”. Neither was true: the six unassigned pads were the retention posts of the DIN 42802 sockets, and there was no M3 hole anywhere in the file. The assembled carrier could not have been fastened into POD-P1 or into the helmet shell. Now: MH1–MH4, 3.2 mm non-plated, at (5, 5), (145, 5), (5, 125) and (145, 125) on the 150 × 130 mm outline of ECO-EEG-018, each with a 6 mm diameter copper keep-out on all four layers. The same pattern is carried into the MP-01 module plate and the POD-P1 bosses. The carrier is held to MP-01 on four M3 × 18 mm nylon hex female-female standoffs with eight M3 × 6 nylon pan screws. Verified: on the fabrication drawing, in the NPTH drill file, and in the DRC keep-out check. Impact: four holes, a copper keep-out, and matching features on two printed parts.

### ECO-EEG-008 – the DevKit would not have fitted

Class: blocking. Found by: checking the socket placement against the ESP32-S3-DevKitC-1 mechanical drawing. Was: J6 at x = 70 mm and J7 at x = 92 mm, a 22.0 mm row spacing. The DevKitC-1 header rows are 22.86 mm (0.900 in) apart. Now: J6 at x = 78.0 and J7 at x = 100.86. Verified: simulate_production.py station 5 measures the spacing from the design source. Impact: placement only.

### ECO-EEG-009 – the firmware pin map used the PSRAM pins

Class: blocking. Found by: checking the pin map against the ESP32-S3-DevKitC-1-N16R8 datasheet. Was: LED_SR_DATA on GPIO35, LED_SR_CLK on GPIO36 and LED_SR_LATCH on GPIO37. On the -N16R8 those three pins carry the octal SPI PSRAM, which the firmware needs for the ring buffer of RFQ F-06. Using them would have made the ring buffer and the shift register mutually exclusive. GPIO45, also used, is the VDD_SPI strapping pin. Now: the shift register moved to GPIO41, GPIO42 and GPIO0, and the microSD interface dropped from four-bit to one-bit SDMMC to release them. The sample stream is 50.7 kB/s of frame payload at 1000 Hz – 1015 bytes every 20 ms – against about 2 MB/s available on one-bit SDMMC, so the headroom is ample. RFQ E-20’s “≈70 kB/s” and F-12’s “≈64 kB/s” are allowances that include STATUS and SIGNATURE frames and filesystem overhead; they remain the requirement and 50.7 kB/s is the payload figure. J7 positions 11, 12, 13 and 15 are explicitly unconnected and are marked so on the assembly drawing and in board_pins.h. Verified: board_pins.h is generated from design.py; the simulator checks that none of the four reserved GPIOs appears in it. Impact: firmware pin map, three carrier nets, RFQ E-18 and E-20 wording. Consequence for J26: GPIO0 is now LED_SR_LATCH, so J26 way 6 is NC_GPIO0, a spare way, and the carrier’s J26 is console and recovery only and cannot enter download mode. End-of-line flashing goes through the DevKitC-1’s own UART USB-C port, which carries the auto-reset circuit on the DevKit itself and is reachable through the MP-01 opening. Ring-buffer size: the “12 MB ring buffer” this ECO cited in Rev A was never possible on an 8 MiB part. See ECO-EEG-025.

### ECO-EEG-010 – four functions were declared and not connected

Class: major. Found by: netlist connectivity check. Was: ENV_CMP (the stimulus comparator of RFQ E-12) had one pad and no comparator; MIC_BIAS had one pad; SPARE1 and SPARE2 (the two spare converter channels of RFQ 3.1, “brought out to the connector”) had one pad each and no protection; and ROOM_PRE had no source at all, because the room microphone that RFQ E-15 requires had no connector. Now: U7 (TLV3201) with its threshold divider, hysteresis and output clamp implements E-12 and it is raised from Should to Mandatory; the electret bias is R89, fitted only if the preamp module does not provide its own (ICD-EEG-006 section 7.2 gives the decision); the two spare channels get protection networks R15/R16, D15/D16, C15/C16 and are brought to J22; J28 is the room-microphone module socket. Verified: every one of those nets is now a connected island; E-12 and E-15 have named test steps in TST-EEG-004. Impact: one new device, one new socket, six new passives, two RFQ requirement changes. Not closed by this ECO: U7’s output logic level is wrong for a 3.3 V input. See ECO-EEG-023, which is open. The EOG panel sockets on J22 are not fitted in a standard build; the sockets, their cable and their drawing are a Phase 2 option and PARTS-EEG-019 lists them with no part number yet.

### ECO-EEG-011 – no analogue reference at the module connectors

Class: major. Found by: review of the module interface. Was: neither analogue module connector carried an analogue reference pin, so the module’s own reference and the carrier’s AGND_REF pour met only through the supply rails. Both modules’ AVDD and AVSS were also hard-commoned, which parallels two regulators. Now: AGND_REF is on J23.4 and J29.4. Module #2’s rails are separate nets (AVDD2, AVSS2) joined to the carrier rails by R92 and R93, which are fitted by default and removed if module #2 regulates its own rails. ICD-EEG-006 section 7.1 gives the measurement that decides; Rev A of this document cited section 4, which is the mechanical stack, and is corrected. Verified: SCH-EEG-005 sheet 4. Impact: two connector pins, two 0 Ω links, one decision recorded in the build record.

### ECO-EEG-012 – the non-plated holes would have been plated

Class: major. Found by: reading the footprint definitions. Was: the six DIN 42802 retention holes were declared (pad "" np_thru_hole circle … (layers *.Cu *.Mask)). A CAM operator taking a naive export gets six plated barrels with mask openings, in the analogue zone, tying the bodies of the three touch-proof EMG sockets to the analogue reference pour. It is a silent defect: the boards look right and fail leakage. Now: non-plated pads carry no copper and no mask on any of the four layers, and both they and the four M3 holes are supplied in a separate -NPTH.drl file. The fabrication drawing calls them out. Verified: the simulator checks that the non-plated hole count matches the design and that no hole appears in both drill files. Impact: footprint definitions, drill output, one fabrication note. Part status: the J15–J17 socket itself is open. design.py names Stäubli SLB1,5-F as a class, not a confirmed PCB part; a touch-proof 1.5 mm socket with a PCB-mount signal pin and two 1.5 mm retention posts must be sourced and first-articled before Phase 2, and AVL-EEG-017 carries a 12-week lead-time risk against it.

### ECO-EEG-013 – the analogue module connectors could not be escaped

Class: blocking. Found by: the router failing to reach the inner row. Was: J2 and J4 were 2×10 sockets carrying twenty analogue nets between them. With no vias permitted under the module outline, the inner row of a 2×10 has no route out: the outer row’s pads block the only direction, and going round the connector needs a corridor wider than the board provides. Now: each is split into a 1×10 analogue-signal socket (J2, J4) and a 1×6 rail socket (J23, J29). Every analogue signal pin now escapes sideways on L1. Verified: every analogue signal pin on J2, J4, J23 and J29 has a route out on L1 in the released routing. What this ECO closes is that the escape is now geometrically possible, which it was not with two 2×10 sockets; it is not a statement about the board as a whole. The full routing result – which now closes, at zero violations and 145 of 145 nets – is stated once, in ECO-EEG-018, and kicad/EEG-CAR-01_RevB_DRC_report.txt is the authority for it. Impact: four connectors instead of two, a change to the jumper schedule in ICD-EEG-006. Note on the move to four layers. Going to four layers does not restore the 2×10 option. Vias are still not permitted under a module outline, so the inner row still has no escape, and the split stands.

### ECO-EEG-014 – the light lines ran through the electrode harness

Class: major. Found by: the router, which could not get eight digital lines from x = 5 mm to the shift register without crossing the whole analogue zone. Was: one 22-way harness socket at the far left of the analogue zone carried twelve electrode conductors – eleven electrode signals and one drain – and ten contact-light conductors, LED1 to LED8 plus LED_V and LED_GND. DSN-EEG-003 Rev A.2 recorded this as finding 10 and accepted it on the grounds that the lights are dark during recording. That was the wrong call: it forces eight digital conductors to run the full width of the analogue zone on the board as well as sharing a cable with eight high-impedance electrode leads. Now: two cables and two sockets. A 12-way screened electrode bundle at J14 in the analogue zone, and a 10-way light ribbon at J30 in the digital zone next to the shift register. Conductor count is unchanged at twenty-two; the crossing is gone. Verified: the DRC checks that no digital net enters the analogue zone and the simulator checks that no digital net shares the electrode harness. Both results are clean in the released routing: the DRC report records zero zone crossings, so no digital net enters the analogue zone. The rest of the routing result, including what is still open in it, is stated once in ECO-EEG-018. Impact: two harness assemblies instead of one (WH-EEG-008), one extra connector, a change to the helmet cable routing in DSN-EEG-002.

### ECO-EEG-015 – part identifiers meant two different things

Class: major, documentation. Found by: cross-document audit. Was: HM-xx was used both as a figure label and as a part number in DSN-EEG-002, and the two namespaces collided. HM-07 named the boom microphone arm in DSN-EEG-002 section 10 and the battery hatch in DSN-EEG-003 section 4, in the STL set, in the kit BOM and in the RFQ scope. A manufacturer printing “HM-07” produced a different part depending on which document was open, and the real boom arm and pod hatch had no file under their own identifiers. Section 10 also requires the part ID to be engraved in the model, so the wrong identifier would have been moulded into the part. Now: PARTS-EEG-019 is the single register. Figures are FIG-nn; HM-xx is reserved for parts; the battery hatch is HM-08 and its file is HM-08_battery_hatch.stl. A migration table maps every v1 filename to its v2 identifier. Impact: one file rename, four documents, the kit BOM and the RFQ scope line. Extended in Rev B: the same failure had been reproduced in the document namespace. Section 1.3 of this document closes it.

### ECO-EEG-017 – lithium shipping was never mentioned

Class: major, compliance. Was: every kit contains an 18650 cell and is posted to a participant and back. Package v1 said nothing about UN3481, packing instruction PI967, state of charge, marking or carrier documentation. Now: RFQ S-09 is new, REG-EEG-012 section 3 states the obligation and PKG-EEG-015 section 7 gives the procedure for both the outbound and the return leg. The procedure has one home, in PKG-EEG-015 section 7; REG-EEG-012 cites it and does not restate it. Verified: the packing list and the shipping label set are checked against PI967 in PKG-EEG-015 section 7. Impact: one new requirement, two documents, the outer carton artwork, and a carrier declaration per shipment.

### ECO-EEG-018 – the routing is now supplied, and what doing it found

Class: scope, and it carries two blocking findings. Renumbered: this change was ECO-EEG-016 in package v2 Rev A, which is also this document’s number. Cite it as ECO-EEG-018. RUL-EEG-021 section B rules the numbering and RUL-EEG-021 section A rules that the board size and the layer count are findings of this work rather than changes in their own right; both are recorded below. The numbers ECO-EEG-028 and ECO-EEG-029, which a Rev B draft briefly used for them, are withdrawn. Was: RFQ-EEG-002A asked a bidder to route the carrier. Now: the routing is supplied, on four layers, and RFQ-EEG-002A becomes a review of it. The routing was produced by the programme’s own constraint-aware router and has not been reviewed by a human layout engineer; the DRC report lists every connection the router had to relax, so a reviewer can go straight to the tight places. The board did not close when this ECO was first written – 122 of 145 nets, 25 violations – and it closes now; section 2B records what was changed to close it and is where the before-and-after belongs.

The routed result. This is the one home for the figures; every other document cites this paragraph or the report behind it, kicad/EEG-CAR-01_RevB_DRC_report.txt, which is the authority. If the two ever differ, the report is right.

| Measured | Value | Against |
| --- | --- | --- |
| Board | 150.0 × 130.0 mm, four layers | – |
| Track segments | 3 745 | – |
| Through vias | 552 | no blind, buried or filled vias |
| Reference planes | one continuous island per net, on both inner layers | the star-point rule of DSN-EEG-003 section 3.3 |
| Smallest clearance, L1 | 0.260 mm | 0.20 mm rule |
| Smallest clearance, the planes | 0.285 mm | 0.20 mm rule |
| Smallest clearance, L4 | 0.275 mm | 0.20 mm rule |
| Narrowest conductor | 0.20 mm | 0.20 mm minimum |
| Connections the router had to relax | 169 – 36 narrower than the 0.25 mm preferred width, 133 at full width with a reduced gap | every one at or above the 0.20 mm minimum conductor and gap |
| Smallest plated hole | 0.30 mm | 0.30 mm rule |
| Digital nets in the analogue zone | zero zone crossings | none permitted |
| AGND_REF-to-DGND bridges | exactly one, at R90 | exactly one |
| HARN_SHIELD-to-DGND bridges | exactly one, at R91 | exactly one |
| Duplicate copper | zero segments, zero via positions | none |
| Nets fully connected | 145 of 145, none without copper | 145 |
| Violations | 0 | zero to release fabrication data |

Nothing in the rule set is breached. The report’s own line is “VIOLATIONS: 0 – none. The board passes every rule listed above”: no clearance violation – the two vias that sat 0.328 mm from an electrode net are gone – and no width, annular-ring, hole-size, board-edge, non-plated-hole, isolation keep-out or via keep-out violation. The isolation keep-out is worth naming, because an earlier routing put inner-plane copper inside it: the strip at x ≥ 141 mm is now free of copper on all four layers, and the report says so. It is free because the keep-out is now enforced against copper edges rather than track centrelines (section 2B). That is the only ground on which the sentence may be repeated. Anyone about to write it in another document reads the report first.

What that does not settle. The board passes every rule it was given, on 169 connections by taking the tightest geometry those rules allow. A board that closes at minimum geometry is not the same board as one that closes at preferred geometry, even when every rule passes; the previous route relaxed 19 connections and this one relaxes 169. Nothing on this board has been manufactured or measured, and no human layout engineer has read the routing.

The gate of section 3 is therefore met and the fabrication data in kicad/ is RELEASED FOR REVIEW under RFQ-EEG-002A. It is NOT RELEASED FOR FABRICATION, and what holds it back is now the review rather than the report: no human layout engineer has read this routing. Reviewing the whole routing – the 169 relaxed connections first – correcting what is wrong and signing it off is the scope of RFQ-EEG-002A. No document in this package may say that the fabrication data is released for fabrication, or that boards may be ordered from it, before that review closes, and no document may state a DRC result without reading the report first, which is RUL-EEG-021 section C ruling 4 applied to the layout.

Verified: tools/emit_all.py regenerates the board, the pours and the DRC report from design.py on every run, and tools/simulate_production.py writes SIM-EEG-018. The simulation run that Rev B pointed at predated the routing close and still carried the DRC violations as an open item. Corrected 2 September 2026: the run behind this release is later than the close – 193 passed, 0 failed, 6 open – and carries no DRC item at all, because the report records no violations. Impact: RFQ section 1.1, section 8, the pricing template, and the covering email.

#### ECO-EEG-018, first finding – the carrier would not close at 130 × 124 mm

Class: blocking. Found by: doing the layout. The router could not close the board at the package v1 size. Was: the carrier was specified as 130.0 × 124.0 mm. Package v1 asserted that size and never routed it. Thirty connectors, 211 parts and 156 nets do not fit: the escape corridors around the module sockets, the 6 mm mounting keep-outs, the isolation keep-out at the right edge and the 0.35 mm electrode-net clearance together leave no channel wide enough to get the analogue zone out to the converters. Now: the carrier is 150.0 × 130.0 mm. The extra 33.8 cm² of bare board costs a few euro per unit at these quantities, against a real risk of an unroutable design; that is not a close trade. Everything dimensioned from the outline moves with it: mounting holes to (5, 5), (145, 5), (5, 125), (145, 125); the zone split to x = 62 mm, analogue left and digital right; the isolation keep-out to x ≥ 141 mm, y = 2 to 22 mm, no copper on any layer. The enclosure grew with the board: POD-P1 base 163.0 × 143.0 × 58.0 mm external and 158.0 × 138.0 × 55.5 mm internal with 2.5 mm walls and a 1.6 × 1.2 mm gasket groove; POD-P1 lid 163.0 × 143.0 × 6.0 mm with a 2.0 mm spigot; MP-01 module plate 146.0 × 126.0 × 3.0 mm with an 8 mm solid border, a field of 12 × 3 mm jumper slots on a 16 × 7 mm grid, Ø2.7 mm M2.5 fixing holes between the slot rows and one 31 × 61 mm opening over the DevKit. The slot figure is a size, not a count: mech_gen.py cuts each slot with .slot2D(12.0, 3.0, 0). The stack budget is floor 2.5 + boss 6.0 + carrier 1.6 + standoff 18.0 + plate 3.0 + modules ≤ 18.0 = 49.1 mm against 55.5 mm internal, a margin of 6.4 mm. Verified: design.py BOARD_W and BOARD_H, the fabrication drawing note 1, the DRC outline and keep-out checks, and mech_gen.py for every printed part. Impact: the board specification in DSN-EEG-003 section 3.2, every mechanical dimension in ICD-EEG-006 section 4, PARTS-EEG-019 sections 2.2 and 2.4, the MECH-EEG-020 drawing set, the foam schedule in PKG-EEG-015, the bare-board price, and the Phase 2 helmet shell, which was sized for a carrier that no longer exists and is recorded as such in RFQ M-01. Consequence for the standoff argument: the slant path from carrier copper, over the edge of the isolation keep-out and up the 18 mm standoff to any host-side conductor on MP-01 is ≥ 18 mm, more than twice the 8 mm the safety case asks for. RISK-EEG-011 SR-08 is closed by that, and the claim in Rev A that the standoff is not dimensioned is withdrawn.

#### ECO-EEG-018, second finding – two layers cannot carry both the routing and the reference

Class: blocking. Found by: doing the layout. This is the single most important thing package v2 learned by doing the work instead of asserting it. Was: two layers, F.Cu and B.Cu. Package v1’s architecture argument was that a two-layer carrier is cheap and easy to route, and package v2 Rev A repeated it. Actually laying the board out showed that it is not true. On two layers the bottom side has to be both the continuous reference plane and the second routing surface, and it cannot be both: every track routed on the bottom cuts the pour that the analogue traces above it are referenced to. What comes out is a swiss-cheesed pour with return currents taking whatever path is left, which is precisely what DSN-EEG-002 section 13’s “layout rules that are requirements, not preferences” forbids. The zoning, star-point and isolation rules that the same layout has to satisfy are DSN-EEG-003 section 3.3. Now: four layers – L1 signal, L2 reference plane, L3 reference plane, L4 signal. That gives two full routing surfaces and a continuous reference under every analogue trace. The reference planes are split the same way on both inner layers: AGND_REF left of x = 62 mm and DGND right of it, on both L2 and L3, tied together by stitching vias, with AGND_REF joining DGND at R90 only and HARN_SHIELD joining DGND at R91 only. Stack-up: mask / 35 µm L1 / prepreg 0.200 / 17 µm L2 / core 1.065 / 17 µm L3 / prepreg 0.200 / 35 µm L4 / mask = 1.60 mm ± 10 %, FR-4 with Tg ≥ 150 °C, outer copper 1 oz and inner copper 0.5 oz. Vias: through vias only. No blind, buried, back-drilled, filled or plugged vias, at 0.60 mm pad on a 0.30 mm finished hole with a 0.15 mm annular ring, tented both sides. The smaller via than package v1’s 0.80/0.40 is what makes the escape corridors fit; through vias only is what keeps the board a class 2 commodity four-layer job rather than an HDI one. Cost: at 2 units a four-layer board is about €35 more in total; at 50 units it is about €3 per board. For a sixteen-channel EEG front end that is the right trade, and it is a real answer to package v1’s cost argument rather than a refusal of it. Verified: design.py fabrication note 2 and the layer map in kicad/gerber/README_layer_map_and_checksums.txt; the pour and stitching checks in the DRC report, which records each reference plane as one continuous island per net; the layer count and stack-up on the fabrication drawing. The counts are in the routing-result table above. Impact: the board specification in DSN-EEG-003 section 3.2, the Gerber layer map and the checksum file, the fabrication drawing, the IPC-6012 class 2 / IPC-A-600 class 2 fabrication scope, the bare-board price line in the RFQ pricing template, the electrical-test netlist, and the argument in DSN-EEG-003 section 2 that package v1 made for two layers, which is withdrawn.

### ECO-EEG-019 – C1–C16 carried an X7R part number under a C0G requirement

Class: major. Found by: cross-document audit, reading the Murata dielectric code in the part number against the requirement in the same table. Was: every document demanded C0G for the electrode-input capacitors and TST-EEG-004 T00 made X7R a hard reject, but the named part was GCM188R71H103KA37D, whose R7 code is X7R. A buyer running the AVL-EEG-017 section 8 checklist line “the order does not contain the word X7R against that line” could not execute it, because the part number itself is an X7R part. The same defect appeared at C21/C41/C61, named as “100 nF C0G 50 V” against GCM188R71H104KA57D, also X7R. Now: C1–C16 are Murata GCM1885C1H103JA16D, 10 nF, C0G, 50 V, 0603.

Corrected 2 September 2026, and the correction changes the answer for C21/C41/C61. Rev B recorded that there was no fix of that kind for the Sallen-Key capacitors, because a 100 nF C0G in 0603 at 50 V is not a stocked part, and specified them as X7R with a stated ±15 % capacitance tolerance over the operating range, Murata GCM188R71E104KA57D, with the C0G requirement holding at C1–C16 and deliberately not met at C21/C41/C61. That stood for one day. The fix was to stop asking for 100 nF: scale the capacitors down by ten and the resistors up by ten, and f₀ and Q land where they were on parts that exist. C21/C41/C61 are now 10 nF C0G – the same GCM1885C1H103JA16D this board already buys sixteen of – C22/C42/C62 are 22 nF C0G GCM1885C1H223JA16D, and R25/R26, R45/R46 and R65/R66 are 215 kΩ. Same 0603 footprints, same nets, no new purchasing line, and the board does not move: it is a BOM change. The only X7R parts left in the envelope channels are C23/C43/C63 and C24/C44/C64, the AVDD and AVSS decoupling at U1–U3, where the dielectric is not an E-11 question; design.py carries them as “100n X7R 25V”, GCM188R71E104KA57D, whose E code is the 25 V the value string states.

Withdrawn on 2 September 2026: the sentence “and the Sallen-Key corner tolerance in TST-EEG-004 is widened to match”. It described a test limit that did not exist. Before T12e was written there was no f₀ limit in TST-EEG-004 to widen, so nothing was widened and the requirement was not made met by this change; TST-EEG-004 section 16 item 16 says so against this ECO by name. What was true was narrower and worse: T12e measures and records f₀ per unit against 42 to 58 Hz, which was the band the fitted X7R parts could hold and is wider than RFQ E-11’s 50 Hz ±10 %, so E-11’s low-pass half was not met with the approved parts and no widened E-11 limit existed to appeal to. Either E-11 had to be restated to that band or a stocked C0G part had to be qualified; at the time neither had been done, and it was carried as RFQ-EEG-001 Rev E section 12 item 15 and TST-EEG-004 section 16 item 16. A dielectric substitution that a test limit is said to accommodate is a substitution nobody has to argue for, which is why the sentence is withdrawn here rather than quietly reworded.

Superseded later the same day: the second of those two alternatives was taken. A stocked C0G part was qualified – the 10 nF the board already buys – and the network was rescaled around it, so E-11’s 50 Hz ± 10 % low-pass half is met in the design: design.py computes f₀ = 49.9 Hz with Q = 0.742 unchanged, and C0G’s ± 5 % with no temperature coefficient holds the corner inside 47.6 to 52.6 Hz against a 45 to 55 Hz band, where the X7R network ranged 42.4 to 57.4 Hz and no build of it could have been held inside. The withdrawal above stands and is not softened by this: no test limit was ever widened, and what closed E-11 was changing the parts, not the limit. T12e’s 42 to 58 Hz band was set around the X7R spread and is now wider than the fitted parts need; whether TST-EEG-004 restates it is TST-EEG-004’s, and this register does not restate a limit it does not own. E-11 is met in the design: nothing has been built and no corner has been measured.

Verified: the part numbers are in design.py and flow to the BOM and the CPL through emit_all.py; AVL-EEG-017 section 1.2 and QP-EEG-010 section 2.3 carry the same numbers. Impact: two AVL lines, the incoming-inspection check, and the BOM. No test limit was widened by this change; T12e’s 42 to 58 Hz band is set by TST-EEG-004 and is a measurement recorded per unit, not an acceptance of E-11.

### ECO-EEG-020 – the board had no fiducials

Class: minor. Found by: assembly review. ASM-EEG-007 section 2.2 recorded that design.py placed no fiducial footprints and proposed a two-point vision teach on the plated pads of TP14 and TP10, priced by the manufacturer. Was: no global fiducial anywhere on the board. Teaching a placement machine on test-point pads is slower, less repeatable and a per-lot cost for no reason. Now: three 1.0 mm round copper fiducials with 3.0 mm mask openings, FID1 to FID3, at (12.0, 10.0), (144.0, 100.0) and (12.0, 120.0) in design.py. The vision-teach workaround is withdrawn. RUL-EEG-021 section B now carries the same three positions. Its first issue transcribed them as (8, 8), (142, 8) and (8, 122); that error is corrected in the register, and under section 1.4 design.py governs in any case. RUL-EEG-021 is a controlled document at Rev A and is registered in section 1.1; the tools/RULINGS.md worksheet it grew out of is not cited. Verified: the footprints are in the placement file and the fiducial layer of the assembly drawing. Impact: three footprints, one assembly-drawing note, and one deleted paragraph in ASM-EEG-007. Panel fiducials remain the fabricator’s choice.

### ECO-EEG-021 – the I²C bus had no pull-ups on the carrier

Class: major. Found by: ICD-EEG-006 section 8 open point 1. Was: SDA and SCL reach J6/J7, J8, J11 and J12 and there was no pull-up resistor anywhere in design.py. The bus depended entirely on whatever pull-ups the purchased modules happened to carry, which is not a design: it changes when a module is substituted, it is not on any drawing, and there is no test that would catch its absence. Now: R94 and R95, 4.7 kΩ 1 %, from SDA and SCL to DVDD3V3, in the digital zone at (132, 26) and (132, 30). Verified: SCH-EEG-005 sheet 5; the nets are one island each in the connectivity check. Impact: two new passives, two AVL lines, one build instruction, and an I²C bus scan in the production test.

### ECO-EEG-022 – VBUS_DET did not reach a guaranteed logic high

Class: major. Found by: ICD-EEG-006 section 8 open point 2 and TST-EEG-004 T21, which recorded the defect and dispositioned it as “design escalation, not rework”. No ECO had been raised against it, so it would have gone to fabrication. Was: R84 = 100 kΩ and R85 = 56 kΩ gives 5 × 56 / 156 = 1.79 V at VBUS = 5.00 V, against a guaranteed V_IH of 2.48 V on a 3.3 V ESP32-S3 input. The first of the two S-01 charge interlocks would not have asserted reliably. Now: R85 changes from 56 kΩ to 150 kΩ, Vishay CRCW0603150KFKEA, giving 5 × 150 / 250 = 3.00 V, comfortably above V_IH. R84 is unchanged at 100 kΩ. Verified: the value is in design.py; TST-EEG-004 T21’s limit band is restated around 3.00 V; the interlock itself is tested end to end. Impact: one resistor value, one AVL line, one test limit band.

### ECO-EEG-023 – ENV_CMP swings ±2.5 V into a 3.3 V pin

Class: major. Status: OPEN. This change is specified and is NOT implemented in design.py, so the defect it describes is still in the design as released. Found by: cross-document audit of the comparator supply nets against the GPIO it drives. Is: U7, a TLV3201, is powered from AVDD (+2.5 V) and AVSS (−2.5 V) and its output CMP_RAW therefore swings ±2.5 V. It reaches GPIO3 through R83 (10 kΩ) and the D23 BAV99 clamp to DGND and DVDD3V3, so the clamp is the only thing between a −2.5 V output and a 3.3 V input, and the logic-high margin is set by the clamp rather than by design. Proposed: power U7 from DVDD3V3 and DGND so that the output swings 0 to 3.3 V into GPIO3 with full margin and D23 becomes a belt-and-braces part. The comparator’s inputs cannot simply follow: AGND_REF sits 2.5 V above DGND, so the inputs must be re-referenced to a DVDD3V3/2 divider with the envelope AC-coupled into it. The TLV3201 has a rail-to-rail input stage that reaches its negative rail, so the topology works, but it moves a signal across the analogue-to-digital reference boundary. Why it is open: this is a change to a circuit that straddles the reference split, and the safety and layout reviewer must check it before it is cut in. Implementing it silently would defeat the review. Impact if implemented: two U7 supply pins, a new divider, one new capacitor, a change to the star-point argument in DSN-EEG-003 section 3.3, and a re-issue of RISK-EEG-011.

### ECO-EEG-024 – S-02’s single-fault DC limit is not met

Class: blocking for Phase 2. Status: IMPLEMENTED 2 September 2026. Rev B carried this ECO as OPEN and not implemented, with 47 kΩ fitted and S-02 not met at 53.2 µA against a 50 µA limit. design.py now fits 68 kΩ at R1–R16 and the calculated single-fault DC patient auxiliary current is 36.8 µA. S-02 is met in the design and is not signed off – see below. Found by: the patient-current budget in RISK-EEG-011 section 4.5, reached independently by WH-EEG-008 section 5.5 and JIG-EEG-009 section 4.3. Was: with R1–R16 at 47 kΩ, the single-fault DC patient auxiliary current is calculated at 53.2 µA against the 50 µA of S-02, which is a 6.4 % overrun and a failure. RISK-EEG-011 classes it Critical. Now: R1–R16 are 68 kΩ, 68k 0.1% 25ppm, Vishay TNPW060368K0BEEA. That gives 2.5 V / 68 kΩ = 36.8 µA, keeps the input corner at 234 Hz, which is −0.75 dB at 100 Hz, and raises Johnson noise to 0.28 µV for a total of 0.31 µV, still well inside E-03. RFQ E-10 moves to its ±1.0 dB at 100 Hz branch, which the requirement already states for exactly this case and which is now the live branch because the 68 kΩ is what is fitted; the ±0.5 dB branch belonged to the 47 kΩ build and is history. Same footprint, same nets, same pad count: the board does not move. Why it was open, and why that reasoning is withdrawn: Rev B held the change back so that the two prototypes could be built with 47 kΩ and the current measured rather than calculated before Phase 2. That left a 6.4 % overshoot on a safety limit standing in the released design through three revisions in order to measure the failing state, when the series resistor is the one component in the patient path whose value can be changed without moving anything else. The measurement does not go away: it changes from measuring a known failure to confirming a calculated pass. What implementing it did NOT close. The safety review. RISK-EEG-011 SR-01 is closed in the design and unsigned; the electrical safety reviewer of RISK-EEG-011 section 7 owns it and that review has not started, which is open item 6 of the simulation run of 2 September 2026. Nothing has been built and nothing has been measured: 36.8 µA is a calculation, exactly as the 53.2 µA before it was. “Met in the design” is what may be written; “met” on its own, and “signed off”, may not. Verified: the value is in design.py and flows to the BOM and the CPL through emit_all.py; tools/simulate_production.py recomputes the single-fault current from it. TST-EEG-004’s patient auxiliary current step, on FIX-04, measures it on the first unit. Impact: sixteen resistor values, the live branch of RFQ E-10, one AVL line, the noise budget in RISK-EEG-011 section 4, and a re-issue of RISK-EEG-011.

### ECO-EEG-025 – F-06’s three-minute ring buffer does not fit the PSRAM

Class: major. Found by: FW-EEG-001 section 5.8, where the allocation returns NULL. Was: RFQ F-06 asked for a ring buffer of at least three minutes at 1000 Hz and put the figure at “≈12 MB in PSRAM”. The mandated module is the ESP32-S3-DevKitC-1-N16R8, which is not substitutable under E-18 and has 8 MiB of PSRAM. Three minutes is 9.13 MB and main.c asked for 12 MiB, so the requirement could not be met on the mandated part and the request failed at run time. ECO-EEG-009 also cited the 12 MB figure in its justification and is corrected above. Now: the ring is 6 MiB, which is 126 seconds of raw samples at 1000 Hz (124 s framed), and RFQ F-06 is relaxed to 90 seconds of ring plus unlimited backfill from the microSD copy. The ring covers a transport dropout; anything longer is recovered from the card, which holds the byte-identical frame stream. The 90-second figure is the requirement and 126 seconds is the fitted margin. Verified: the constant is in the firmware build and reported in the STATUS frame; TST-EEG-004’s dropout-and-recovery step exercises the backfill, not just the ring. Impact: one RFQ requirement, one firmware constant, the IFU sentence that told the participant the instrument keeps three minutes, and the test step that checked it.

### ECO-EEG-026 – E-04’s crosstalk limit is neither achievable nor measurable

Class: major. Found by: ICD-EEG-006 section 3.2 and raised at ICD-EEG-006 section 8 item 7 without ever reaching the RFQ. Was: RFQ E-04 asked for −100 dB channel-to-channel crosstalk. That is not achievable through a 60 mm un-interleaved ribbon, and it is about 40 dB below this instrument’s own noise floor, so it is not measurable on this hardware either. TST-EEG-004 T9a nevertheless tested to −100 dB with 6.2 dB of detection-floor margin, and measured only the carrier. Now: E-04 is restated as −80 dB at 50 Hz, measured on the carrier, with the ribbon’s contribution characterised once on the first prototype and recorded rather than tested per unit. The restatement is honest about what the instrument can see: a limit that cannot be measured is not a limit. Verified: TST-EEG-004’s crosstalk step is rewritten to the −80 dB carrier measurement and JIG-EEG-009’s detection-floor calculation is restated against it. Impact: one RFQ requirement, one test step, one fixture note, and one characterisation report on the first prototype.

### ECO-EEG-027 – the envelope AC-coupling corner removed the envelope

Class: major. Found by: ICD-EEG-006 section 8 item 4, computing the corner from the fitted values. Was: C20/C40/C60 were 1 µF into R20/R40/R60 at 10 kΩ, which is a 15.9 Hz corner, not the 0.1 Hz of RFQ E-11. design.py’s own descriptive text claimed 0.1 Hz and AVL-EEG-017 listed the part as “1 µF … 0.1 Hz corner”. For the envelope of a speech signal 15.9 Hz is not merely the wrong number, it is the wrong side of the signal: it removes the envelope the stage exists to pass. Now: C20/C40/C60 change from 1 µF to 10 µF, Murata GRM188R61C106MA73D, with R20/R40/R60 unchanged at 10 kΩ, giving a 1.6 Hz corner. RFQ E-11’s AC-coupling clause is restated as ≤ 2 Hz, because 0.1 Hz into 10 kΩ needs 160 µF, which is not a 0603 part and buys nothing for a speech envelope. Verified: the values are in design.py; the corner is computed in simulate_production.py station 9; TST-EEG-004 measures it and the deviation note that recorded 15.9 Hz is deleted. Impact: three capacitor values, one AVL line, one RFQ requirement, one test limit, and one corrected description in design.py.

### ECO-EEG-030 – an external layout engineer read Rev B, and Rev B is withdrawn

Class: scope, with seven findings. Found by: Danny Van den Bliek, DVDB-electronics BV (Belgium), reading kicad/EEG-CAR-01_RevB_routed.kicad_pcb between 3 and 5 September 2026 under RFQ-EEG-002A.

Was: the Rev B routing was RELEASED FOR REVIEW under RFQ-EEG-002A and was waiting on a human layout engineer, which is the sixth open item of the production simulation and the fabrication-release gate of section 3. Nothing else was thought to stand between the data and a fabrication order.

Now: the reviewer declined the paid review and advised a full redesign from the schematic. The programme accepted that on 4 September 2026. Rev B is withdrawn from fabrication: no board is ordered from it, and Dekimo and DVC were told so on 3 and 4 September 2026. The Rev B data stays in the tree under its own file names as history and is not deleted; it is no longer the current fabrication set.

The seven findings. They are geometric, not circuit errors. Six are rules a design-rule check can enforce and the programme’s own router was never given; the seventh is a missing check rather than a missing rule.

| # | Finding | What it is |
| --- | --- | --- |
| 1 | drills in pads | vias placed inside SMD pads |
| 2 | net stubs and loops | dangling track ends; more than one copper path between the same two pads on one net |
| 3 | sharp inside corners | an angle below 90 degrees between two consecutive segments of one track |
| 4 | off-centre pad entry | a track entering a pad other than through the pad centre along the pad’s major axis |
| 5 | under-width critical nets | no net classes: power, analogue input, reference and light-drive nets routed at the generic width |
| 6 | no module outlines for collision checking | no 3D bodies for the thirteen module assemblies, MP-01 or the DevKit, so no placement can be shown to be free of interference |
| 7 | schematic not logically drawn | no native schematic exists; the eight-sheet PDF is drawn by tools/schematic.py and is not a netlist-bearing schematic a reviewer can read or import |

The reviewer’s stated principle, which the programme has adopted: placement drives via count, must precede routing, and must be reviewed before routing starts.

What Rev C is. Rev C is a board revision that changes footprints, part numbers and the layout. It changes no pin assignment, no net name, no connector coordinate, no outline, no stack-up, no zone split, no keep-out and no star point; if a later change finds one of those necessary it is its own ECO and forces a regeneration and diff of firmware/main/board_pins.h. Rev C is produced in-house as a reviewable design and not as a fabricable one: the seven findings are encoded as rules and checks, a native KiCad schematic and 3D bodies are added, and placement and routing are bought from the JLCPCB layout desk. tools/router.py is no longer the producer of a fabrication set. Its two remaining roles are to hold and enforce the rules so that a returned board can be graded against them, and, optionally, to produce an in-house reference route for comparison.

RFQ-EEG-002A is re-scoped. RFQ-EEG-001 Rev E section 1.1 defines 002A as the review and sign-off of the supplied carrier routing. There is no supplied carrier routing to review any more. 002A is now the review of the layout contractor’s PLACEMENT of Rev C, taken at the placement-confirmation gate and before routing is confirmed, as a written findings list by net, pad pair and coordinate rather than as edits to the board file. It is capped at four to six hours and is agreed in principle with the same reviewer. The full-redesign quote of EUR 3,200 (estimated EUR 3,200 to 4,500 plus VAT) is logged as a grant budget line and is not ordered; the envelope for the current phase is about EUR 500.

Two revision letters in design.py. REV stays at "B" and keeps its file names; REV_C is new and is "C", and stem() computes the file stem from a letter so that no Rev C file name is typed by hand. Rev B artifacts are unchanged in name and content by this ECO.

Verified: tools/emit_all.py --cached and tools/simulate_production.py were re-run on the tree before any change. The DRC report reproduces byte for byte: 0 violations, 145 of 145 nets connected, 3 745 segments, 552 vias, minimum clearances 0.260 / 0.285 / 0.285 / 0.275 mm. The withdrawal is a change of release state, not of geometry, and it adds no violation, opens no net and puts no copper in the isolation keep-out.

Impact: the board – none, Rev B’s geometry is untouched and Rev C has no routing yet. The firmware – none, no pin assignment moves. The mechanical parts – none. The BOM – none in this ECO; ECO-EEG-031 carries the BOM corrections. The test specification – none; the fixture data of JIG-EEG-009 is built against connector positions, and those do not move. The safety case – none: no electrode-path, isolation, battery, charge-interlock or patient-current value changes, so RISK-EEG-011 is not re-issued for this ECO. Units already built – none exist.

### ECO-EEG-031 – four part numbers, and only three of them were known

Class: major. Found by: Bittele/7pcb, who reported the first three to the programme and to Dekimo on 3 September 2026; the fourth by tools/footprint_audit.py, written for this ECO, on 8 September 2026.

Was: the carrier BOM carried four lines that could not be bought as written or could not be fitted to the land pattern beside them, and one class of part that a buyer could substitute without knowing it must not be.

Now, item by item.

1. U1, U2, U3 – TI OPA4376AIDR on a SOIC-14 land. The quad OPA4376 is offered in one package: TSSOP-14, TI’s PW. The D suffix is SOIC and there is no SOIC-14 quad, so OPA4376AIDR is not an orderable part number. The part is OPA4376AIPWR and the footprint is TSSOP-14_4.4x5mm_P0.65mm. What makes this worse than a typing error is the record: Rev A of AVL-EEG-017 quoted OPA4376AIPWR, which is correct, and its Rev B withdrew it on the ground that it did not fit the SOIC-14 land pattern design.py placed – the package was corrected to match the footprint instead of the footprint to match the package. That withdrawal is withdrawn. fplib’s SOIC-14 docstring argued the choice on the two-layer escape between adjacent lands; the board went to four layers at ECO-EEG-018 and the argument had already lapsed, but the deciding fact is simply that the package does not exist. Pin numbering is identical between SOIC-14 and TSSOP-14, so no net moves.

2. L1 – Murata BLM18PG601SN1D, 600 Ω at 100 MHz, 1.5 A. The reviewer read the impedance code as some other impedance; that reading is wrong. Murata’s part-numbering page states that the three impedance figures are two significant digits and a count of zeros, so 601 is 600 Ω at 100 MHz and the design intent is correctly encoded. The defect is the series: PG is the large-current power-supply family, and the same catalogue’s impedance map does not offer it at 600 Ω in 1608 – the 0603 P/S range stops at 470 Ω (1 A). There is no such part. 600 Ω at 100 MHz in 1608 is a general-series part: BLM18AG601SN1D, 600 Ω ±25 %, 500 mA rated, 0.38 Ω DC resistance. The “1.5 A” came with the wrong part number and has no requirement behind it; L1 feeds VDD_ISO, the ADuM4160 module’s device-side supply, which is tens of milliamperes, so 0.38 Ω costs about 10 mV.

3. J15, J16, J17 – Stäubli SLB1,5-F / LB-I1,5. The part number is a class and not a qualified PCB part, which AVL-EEG-017 §1.4.1 and ICD-EEG-006 §1 both already said. What was missing is that a buyer reading the BOM could not see it. design.py now carries a NOT_SUBSTITUTABLE table – AVL-EEG-017 §6.4 remains the home of the list – and the generated BOM prints the reason on the line: PATIENT-CONNECTED under IEC 60601-1, approved in writing from a sample, no “or equivalent”. The same table marks R1–R16, C1–C16, D1–D16, R90 and R91. tools/footprint_audit.py reports J15–J17 as OPEN, never as passes.

4. SW1, SW2, SW3 – Omron B3F-4055. This one was not on anybody’s list. Omron splits the B3F family by body size: 6 × 6 mm is B3F-1000/-3000/-6000, and 12 × 12 mm is B3F-4000/-5000/-5001. B3F-4055 is a 12 × 12 mm switch. The footprint beside it is the 6 × 6 mm four-terminal pattern on 6.5 × 4.5 mm centres, the three switches sit on a 14 mm pitch, and AVL-EEG-017 §1.4’s own row asks for a 6.0 × 6.0 mm body at 160 gf ±50 – which B3F-4055, at 260 gf, does not meet either. The part is B3F-1052: 6 × 6 mm, projected plunger 7.3 mm, 1.47 N {150 gf}, four terminals, no ground terminal. Projected, because a 12 mm B32-series key top mounts only to a projected plunger and the pod lid’s 12.4 mm opening is cut for that cap. The footprint’s own name was wrong with it – SW_PUSH_6mm_H5mm, “5 mm actuator” – and is now SW_PUSH_6mm_H7.3mm; the height does not change the land pattern, it changes the collision model of ECO-EEG-033. RFQ-EEG-001 Rev E E-26 still reads “a 6 mm tactile switch (Omron B3F-4055 class)”, a requirement that contradicts itself inside one sentence. It is not re-issued here; it is recorded as an open item for the next issue of that document.

A fifth thing this ECO changes, and it is not a part number. tools/fplib.py’s TSSOP-14 land pattern was 1.45 × 0.45 mm pads at x = ±2.85 with a 7.50 × 5.70 mm courtyard. fplib’s stated contract is that a house re-importing the netlist into its own CAD gets the same land pattern as KiCad’s standard library, and KiCad’s TSSOP-14_4.4x5mm_P0.65mm is 1.475 × 0.400 mm at ±2.8625 with a 7.70 × 5.50 mm courtyard. Nothing was placed on the old figures, so this costs nothing; leaving them would have put U1–U3 on a pattern that is nobody’s.

Three drawing notes had become false and are corrected in the same edit, because design.py’s NOTES are printed onto the fabrication and assembly drawings: the series resistor was still described as 47 kΩ where ECO-EEG-024 made it 68 kΩ on 2 September; the analogue rails were still said to arrive at “J2 pins 11–14”, and J2 has ten ways – they arrive at J23, the 1×6 rail socket, which is what ECO-EEG-013 split out; and the isolation and mounting keep-outs were still stated as clear “on both layers” on a board that has had four since ECO-EEG-018.

Verified. tools/footprint_audit.py audits all 211 designators against the package, way count and, where the part number encodes it, the value, tolerance, dielectric and voltage: 183 PASS, 0 FAIL, 3 OPEN (J15–J17), 0 UNRECOGNISED, 25 fabricated features. The report is docs/footprint_audit_RevC.md and it is generated, not written. pcbgen.BoardV2().validate() returns 0 errors on 211 designators, 636 pads and 156 nets – no courtyard overlap, no edge breach – with the TSSOP courtyards, which are shorter than the SOIC ones they replace. The netlist is unchanged and that is measured, not asserted: the pad-to-net mapping from design.py was compared against the released kicad/gerber/EEG-CAR-01-IPC-D-356A.ipc excluding its via records, and the two are identical across all 156 nets.

Impact. The board – U1–U3’s land pattern and courtyard change and their placement lines in DSN-EEG-003 §3.3 rule 3 are unaffected, because the protection networks are the R–D–C rows and not the op-amps; nothing else moves. The firmware – none; no pin assignment changes. The mechanical parts – the switch is 7.3 mm to the top of the plunger where the drawing notes said 5.0 mm, which is a collision-model input for ECO-EEG-033 and does not move a printed part; the pod lid opening at 12.4 mm is unchanged. The BOM – four lines. The test specification – none. The safety case – none of the four changes alters the electrode path, the isolation, the battery, the charge interlock or the patient-current budget, so RISK-EEG-011 is not re-issued; J15–J17’s marking makes an existing safety property visible to a buyer rather than changing it. Units already built – none exist.

One consequence for the tools. emit_all.py no longer regenerates the Rev B set. Rev B’s routing was produced against SOIC-14 pads; running the old path from this source would write Rev B file names with Rev C pads and produce a set whose copper does not reach its own pads. The Rev B fabrication set is frozen, is recovered from git rather than regenerated, and emit_all.py --rev-b refuses and says so.

### ECO-EEG-032 – the rule set the DRC was measuring against was incomplete

Class: major. Found by: ECO-EEG-030. This ECO is that ECO’s findings 1 to 5 turned into checks, and it is raised separately because it changes the tools rather than the design.

Was: kicad/EEG-CAR-01_RevB_DRC_report.txt says “VIOLATIONS: 0 – none. The board passes every rule listed above”, and it is right. The words that matter are “listed above”. The rule set had twelve rules and none of them looked for a via inside a pad, a dangling track end, a redundant copper loop, an acute corner, an off-centre pad entry or a conductor narrower than its own net class needs – and the last of those could not be looked for, because there were no net classes with minima to look against. A report can be true and useless at the same time, and this one was: it was the evidence the board was ready for review, and the reviewer declined the review.

Now: the rule set lives in tools/rules.py as data, and it is written down once.

- Net classes. Eight, replacing Rev B’s six: ELECTRODE, ANALOGUE_REF, POWER, USB, ANALOGUE, LED_DRIVE, DIGITAL, DEFAULT. Every one of the 156 nets falls in exactly one. Each class carries a minimum and a preferred width and clearance, the layers it may use, and whether it may have a via at all. Rev B’s table carried one width and one clearance per class and no minimum, which is why all 169 relaxed connections were free to fall to the 0.20 mm board floor with nothing to say that some of them were the wrong nets to relax. Rev B’s figures survive as the preferred column.

- Where the minima come from. POWER at 0.40 mm because the worst case is about 610 mA at J13 and IPC-2221 gives 0.40 mm of 35 µm outer copper about 1.29 A at a 10 °C rise. ANALOGUE_REF at 0.30 mm for source impedance, not current – the rails carry about 10 mA. LED_DRIVE at 0.25 mm on 1.3 mA a site. ELECTRODE at 0.25 mm, above the board floor, because a patient-connected conductor is the last place to spend the last 0.05 mm. Four numbers are proposals and not transcriptions and each is flagged confirm in the source and printed as such in LAY-EEG-034 section 5: the ELECTRODE minimum width, the ELECTRODE layer and via restriction, the ANALOGUE_REF minimum width and the pad-entry tolerance.

- ELECTRODE is L1 only and may not have a via. That is DSN-EEG-003 section 3.3 rule 3 read literally – these nets run on L1 with the reference plane continuous beneath them, and a via breaks the plane it is routed over. It constrains the contractor and it is achievable: J14, the R/D/C protection rows and J2/J4/J23/J29 are all on the top side.

- Six checks in the in-house DRC, in tools/drc_geometry.py, each with its own counted line in the report: via inside an SMD pad; dangling track end; redundant copper path outside a plane; angle below 90° between two segments and any segment off a 0/45/90/135° axis; a track entering a pad more than 25 % of the pad’s minor dimension off its centre; and per class, width, layer and via rule.

- The same rules exported to KiCad. kicad/EEG-CAR-01_RevC.kicad_dru is written from rules.py and is 37 rules. rules.check_dru() balances its brackets on every emit, because a .kicad_dru with one bracket missing loads as far as the error and silently drops every rule after it – which is this whole ECO’s failure mode wearing a different hat.

What KiCad cannot check, said plainly rather than implied. The angle rule and the pad-entry rule have no KiCad constraint and are not expressible in a .kicad_dru. They are checked by tools/drc.py on the geometry that comes back, and LAY-EEG-034 section 6 says so against each of them rather than leaving a reader to assume the contractor’s DRC will catch them. Dangling ends and redundant loops are KiCad built-in checks and are raised to errors in the project file rather than restated as custom rules.

Verified, and this is the evidence that the findings were real. tools/grade_revb.py regrades the released Rev B geometry under the new rules – the pads, tracks and vias out of kicad/EEG-CAR-01_RevB_routed.kicad_pcb and the plane polygons out of tools/routed.pkl, checked against the released census of 3 745 segments and 552 vias before it grades. It reads the released artefact and not design.py, because design.py has moved to Rev C pads.

| Finding | Rule | Occurrences in the released Rev B routing |
| --- | --- | --- |
| 1 | vias inside an SMD pad | 67 |
| 2a | dangling track ends | 5 |
| 2b | redundant copper loops outside a plane | 221, on 22 nets |
| 3 | angle below 90° between two segments | 171 |
| 3 | segments off a 0/45/90/135° axis | 158 |
| 4 | off-centre pad entries | 299 |
| 5 | conductors under their class minimum | 322 |
| 5 | net classes on a forbidden layer | 50 |
| 5 | vias on a class that forbids them | 18 |
|  | total | 1 311 |

The loop line counts loops and the detail lists one line per net, so those two figures differ on purpose; every other line is one occurrence per detail line. The report is kicad/EEG-CAR-01_RevB_regraded_ECO-EEG-032.txt.

None of this is a regression and none of it is reverted. Section 3’s ECO gate forbids a change that adds a violation; a change that adds a check is not in that class, and the paragraph now says so. The Rev B DRC report is not withdrawn either – it is correct about the twelve rules it had, and it is the report on a set that is withdrawn from fabrication for other reasons.

Two of the 1 311 are worth naming, because they are in the register already. The 67 via-in-pad occurrences include the stub_via at U1.3 that design.py’s TARGETED_REPAIRS records as a deliberate repair of 2 September 2026: a 1.0-1.5 mm stub from the op-amp’s AGND_REF pin to a through via. It closed a net and it is finding 1. The 322 under-width conductors are not the 36 of the “169 relaxed connections” restated: that figure counted connections against the board’s single 0.25 mm preferred width, and this one counts segments against eight per-class minima.

Impact. The board – none; no geometry moves and Rev C has no routing to grade. The firmware, the mechanical parts, the BOM, the test specification – none. The safety case – none directly, though the ELECTRODE class now states in machine-readable form what DSN-EEG-003 section 3.3 rule 3 states in prose, which is a strengthening of an existing control and not a new one. Units already built – none exist.

One structural change. design.py’s NETCLASS table and netclass_of() are removed and their callers now read rules.py. design.py keeps the net-name membership sets, which are circuit facts; rules.py holds the geometry, because it has to be exported to KiCad as well as measured, and a rule kept in two places is a rule that will disagree with itself. tools/kicad_parse.py gained segment and via readers so that a released board file can be graded without its source.

### ECO-EEG-033 – the Rev C reviewable design set

Class: scope. Found by: ECO-EEG-030, findings 6 and 7, and the input list the programme committed to the layout contractor on 8 September 2026.

Was: four of the six inputs promised to the layout desk did not exist in any form – a native schematic, a board outline and fixed connector DXF, 3D bodies of the thirteen module assemblies, and a rule sheet – and a fifth, the netlist, existed only as IPC-D-356A.

Now, and this entry grows as each part lands.

1. The native schematic (finding 7). kicad/EEG-CAR-01_RevC.kicad_sch and nine sheets under it, written by tools/emit_kicad_sch.py from the same design.py data the eight-sheet PDF is drawn from. Ten files, KiCad 8 format, hierarchical, one sheet per block: power, isolation and USB, controller, the two analogue front ends, the harness and its sixteen protection networks, audio and the envelopes, the contact lights, and the star points and test points. All 211 designators are placed and the assignment is checked to be a partition of design.C – a designator on two sheets or on none fails the build.

How connectivity is expressed. Power and ground are global through power symbols and carry no sheet pin. Every other net with pads on more than one sheet – 46 of the 156 – is a hierarchical label on each sheet that uses it and a sheet pin on the root. The remaining 110 are local to one sheet or global through a power symbol.

What makes it readable rather than merely correct, which is what finding 7 is about. The quad OPA4376 is drawn as five units – four amplifiers and a power unit – and not as a fourteen-pin box. Each of the sixteen protection networks is drawn as the straight line DSN-EEG-003 section 3.3 rule 3 asks the layout to place it in: resistor, then clamp, then filter capacitor, from the harness towards the module. Each envelope channel is drawn along its signal path. Notes wrap, and a sheet whose content runs off its page fails the build rather than emitting a PDF with the comparator missing.

The gate, and it is a measurement and not an opinion. tools/sch_netlist.py reads the emitted files back – pin geometry out of their own embedded lib_symbols, wires, junctions, labels, hierarchical labels and power symbols – resolves the connectivity geometrically and diffs it against design.py. It does not ask the generator what it meant.

|  | design.py | schematic |
| --- | --- | --- |
| nets | 156 | 156 |
| pins | 614 | 614 |
| structural problems | – | 0 |
| netlist differences | – | 0 |

It earned its keep on the first run: four pull-ups – R50, R51, R52 and R94 – were wired by geometry rather than by pin number and had pin 2 where design.py has pin 1. A resistor is symmetric so the circuit was right and the netlist was wrong, which is precisely the class of error that survives a human reading a schematic and does not survive this check. R94 and R95 were also drawn as a divider and are two separate I2C pull-ups. Both are fixed and the helper that caused it now wires by pin number.

ERC. KiCad is not installed on the machine this was generated on, so kicad-cli sch erc was NOT RUN and no ERC result is claimed. What was run is the structural half: every pin of every placed symbol is on a net, no designator is placed twice and every unit of a multi-unit part is placed exactly once, no net has one pin unless design.py declares it as one of the eleven single-pad nets, and every power net is driven by a power symbol. Zero findings.

The symbol library. kicad/EEG-CAR-01.kicad_sym, generated. Rev B of ICD-EEG-006 said no such file existed and that its own Rev A was wrong to claim one; that was true of board Rev B and is superseded here, and the ICD is corrected in the same change rather than left contradicting the tree. The contractor needs it: a schematic whose symbols live only inside itself opens, and nobody can edit it.

2. The unrouted board, the project and the DXF (inputs 2 and 4). kicad/EEG-CAR-01_RevC.kicad_pcb, written by tools/kicad_pcb8.py at (kicad_pcb (version 20240108)). tools/kicad_write.py, which writes the Rev B routed board in the KiCad 5 dialect, is untouched: that dialect was the right choice for a fabrication set going to houses on older tools and is the wrong one here, because net classes, rule areas and custom rules do not survive it and they are the whole point.

What is in it: the outline on Edge.Cuts and the four-layer stack-up of DSN-EEG-003 section 3.2 in (setup); all 211 footprints with their pads and nets and no track, no via and no arc; AGND_REF and DGND zones on In1.Cu and In2.Cu split at x = 62 mm; and rule areas under exactly the names the .kicad_dru refers to – ISOLATION_KEEPOUT on all four copper layers, ANALOGUE_ZONE, DIGITAL_ZONE, four MOUNT_KEEPOUT_MH* and NO_VIA_J2/J4/J23/J29.

37 footprints are locked: J1 to J30, MH1 to MH4 and FID1 to FID3. The connectors are fixed by the pod, the module plate and the harness and the contractor may not move them; everything else is at its Rev B coordinate, unlocked, and is theirs to place.

kicad/EEG-CAR-01_RevC.kicad_pro carries the eight net classes, an assignment for all 156 nets, the board minima, and the rule severities that make track_dangling an error – which is finding 2a handed to KiCad’s own built-in check rather than restated as a custom rule.

The board is EEG-CAR-01_RevC.kicad_pcb and not EEG-CAR-01_RevC_unrouted.kicad_pcb, and that is deliberate. A KiCad project expects its board and its schematic to share the project’s stem. A differently named board opens standalone – which means without the project file, so the net classes, the assignments and the severities would all be silently absent for the contractor. That the board is unrouted is said in its title block, in its comments, in LAY-EEG-034 and in the handover README; a file name does not have to repeat it, and a broken project association costs more than the name is worth.

kicad/EEG-CAR-01_RevC_outline_and_fixed_connectors.dxf is the mechanical input: outline, MH1-MH4 with their 6 mm keep-outs, the isolation strip, the zone split and all thirty connector courtyards with pin-1 marks, in design.py’s convention – top-left origin, Y down – which is the board file’s convention too, so the two overlay without transforming anything. The CAM convention is the other one and the DXF says so on its own face as well as in the README.

3. The 3D bodies and the collision check (finding 6). tools/step_write.py, tools/mech_bodies.py and tools/collision_check.py.

cadquery could not be installed on the machine this was generated on – 472 MB free where cadquery-ocp needs several times that – so tools/mech_gen.py cannot even be imported there, let alone run, and no printed part was regenerated. Rather than claim a body set that does not exist, step_write.py writes the bodies the check needs from first principles with no dependency at all: FACETED_BREP, planar faces, one solid per extruded polygon. It cannot do fillets, holes in a face or curved surfaces, and does not pretend to – a standoff is a hexagonal prism and MP-01 is four rectangles around its DevKit opening. No STEP reader was available to verify the files; they are checked structurally, and that is what is claimed.

Written: one body per footprint class, 22 of them, bound into the .kicad_pcb so the contractor’s 3D view is populated without importing anything; the carrier with the height envelope of all 211 parts; MP-01 with the four M3 x 18 standoffs; and the thirteen module assemblies.

The twelve module envelopes are DECLARED MAXIMA and no vendor drawing was consulted for any of them. They are registered in ASM-EEG-023 under MECH-D6-MODULE-ENVELOPES, which is new and blocks a build. The placement of the modules on MP-01 is a proposal for the same reason: ICD-EEG-006 section 4 deliberately does not drill the plate to a pattern.

What the check found, and three of the four are open.

| Case | Result |
| --- | --- |
| A – carrier parts against the plate underside | clear. The tallest carrier part is J15 at 10.00 mm against a plate underside at 18.0 mm: 8.00 mm of clearance |
| B – the DevKit against the plate opening | THE OPENING IS 2 mm TOO SHORT. J6 and J7 are 1x22 sockets, so the pin span is 53.34 mm, and the ESP32-S3-DevKitC-1 PCB is 63.0 mm long: it needs 3.2 to 66.2 mm in carrier Y and MP-01’s opening runs 4.0 to 65.0 mm. The board fouls the plate at both ends |
| C – modules against each other and their jumper reach | six of twelve cannot be placed at all, and four cannot reach their connector within the 60 mm of ICD-EEG-006 section 3.2 from ANY position on the plate – M4 to J10 is 86 mm at best, M5 to J11 is 72 mm at best. That 60 mm is a crosstalk limit, not a mechanical one |
| D – plan area, which nothing in this package had ever checked | the twelve envelopes need 17 356 mm2 and the plate offers 12 409 mm2: 140 % fill |

Case D is the one that matters and it does not rest on the assumptions. The two ADS1299 boards alone are 7 345 mm2, 59 % of the net usable plate, and that figure is the Raspberry Pi HAT standard rather than a guess – the PiEEG-8 is a Pi shield. Halving every other envelope still leaves 12 351 mm2 against 12 409 mm2. ICD-EEG-006 section 4’s stack budget finds 6.4 mm of margin in HEIGHT and says nothing about area, and area is the binding constraint.

Nothing here is closed by this ECO. Four routes would each move the answer – a larger plate, a second tier of standoffs, an ADS1299 breakout that is not a Pi HAT (AVL-EEG-017 section 2 M1 permits the breakout to change under the section 6 qualification even though the DEVICE may not), or moving the modules whose jumpers cannot reach onto the carrier as fitted parts. None is decided and none is costed. They are for the programme.

4. The rule sheet (input 6). LAY-EEG-034 Rev A, registered in section 1.1 and issued 8 September 2026. Eight sections: scope and precedence, board and stack-up, zoning and star points, the isolation keep-out, the net-class table, the seven geometry rules, placement, and the deliverables expected back.

It is a controlled document and it is generated, which is not a contradiction: the tables are transcribed from tools/rules.py by tools/emit_rule_sheet.py, so the sheet the contractor signs against and the .kicad_dru the contractor’s DRC loads come from one source and cannot drift. The prose around them is written.

Two things in it are said plainly rather than left to be assumed. Section 6 says which of the seven rules KiCad can check and which it cannot – the angle rule and the pad-entry rule have no KiCad constraint, and the sheet says against each of them that this programme checks it on the returned geometry with tools/drc.py, naming the tool and pointing at the 1 311 occurrences it found in Rev B. And section 7 carries the open item from part 3 – that the modules do not fit MP-01 and that the plate’s DevKit opening is 2 mm short – because a placement reviewer will meet it, while saying that neither is a carrier layout question and neither blocks the order.

The brief for this work asked for one page. It is three. The net-class table is eight rows with a derivation for each, and the geometry table is seven rows with a checked-by column; compressing either into a page would mean dropping the derivations, and a rule a contractor cannot see the reason for is a rule that gets relaxed at the first difficulty. The sheet is still one document and is read in one sitting.

5. The handover set. kicad/RevC_layout_inputs/, written by tools/emit_handover.py: 51 files, 2.6 MB – the project, the ten schematic sheets, the symbol library, the board, the .kicad_dru, the netlist, the BOM and the footprint audit, the two provisional CPLs, the DXF, twenty-five STEP bodies, the collision report, the rule sheet in both forms, a SHA-256 for every file, and README_layout_inputs.md.

The README is the covering note and it says the four things a contractor otherwise has to guess: open the .kicad_pro and not the board, because a board opened alone has none of the net classes or the severities; the KiCad version, with an offer to re-emit for another one; what is locked (37 footprints) and what is theirs (174), and that the unlocked coordinates are Rev B’s and are a starting point rather than a specification; and which two rules KiCad cannot check and that this programme checks those itself on return, naming the tools. It also names the placement gate and asks that routing not begin on an unconfirmed placement.

Nothing in tools/ sends anything. The directory is written; it is uploaded by hand.

The build is deterministic and that is measured, not asserted: two consecutive runs of tools/emit_all.py produce byte-identical checksums across all 51 files. So a checksum that does not match means a file changed, which is the only thing a manifest is worth.

ECO-EEG-033 closes here. What it did NOT do is also part of the record: no ERC and no DRC was run on any of it, because KiCad is not installed on the machine that generated it, and no printed part was regenerated, because cadquery could not be installed there either.

Verified. The board is parsed back after writing and checked against design.py: 211 footprints, 636 pads, 156 nets, 0 segments, 0 vias, 37 locked, netlist identical, every rule area present under the name the .kicad_dru uses, the project file’s classes and its 156 assignments complete, and track_dangling at error. kicad-cli pcb drc was NOT RUN: KiCad is not installed on this machine. On an unrouted board its only finding would be unconnected items, one per net, and that is not a substitute for the check.

### 2.2 Requirement changes and where they come from

Every RFQ-EEG-001 requirement that changed between Rev C and this release, with the change that caused it. The requirement text lives in RFQ-EEG-001; this table is the index, not a restatement.

| Requirement | Change | Source |
| --- | --- | --- |
| E-04 | crosstalk restated from −100 dB to −80 dB at 50 Hz, measured on the carrier | ECO-EEG-026 |
| E-10 | ±1.0 dB at 100 Hz. That is the branch E-10 already carried for this case, and it became the live one on 2 September 2026 when ECO-EEG-024 was implemented and R1–R16 went to 68 kΩ. The ±0.5 dB branch belonged to the 47 kΩ build and is history, not an alternative | ECO-EEG-024, implemented |
| E-11 | AC-coupling corner restated from 0.1 Hz to ≤ 2 Hz, and met at a calculated 1.6 Hz. The 50 Hz ± 10 % low-pass is unchanged and, since 2 September 2026, met in the design: the Sallen-Key moved X7R → C0G at 10 nF / 22 nF with 215 kΩ, giving f₀ = 49.9 Hz and a corner inside 47.6 to 52.6 Hz against the 45 to 55 Hz band. It was not met with the X7R parts, which ranged 42.4 to 57.4 Hz. Unchanged here still means not restated | ECO-EEG-027 for the AC coupling; the low-pass closed under ECO-EEG-019, 2 September 2026 |
| E-12 | raised from Should to Mandatory, because the comparator is now fitted | ECO-EEG-010 |
| E-18, E-20 | pin map and throughput wording; 50.7 kB/s of payload stated against the ≈70 kB/s allowance | ECO-EEG-009 |
| E-23 | rewritten as a separate charge-only receptacle, and the 45 °C charge inhibit restored to E-23 rather than cited from S-04 | ECO-EEG-003 |
| E-24 | live non-conformance: the named isolator module presents USB-B against a USB-C requirement; interim answer is the WH-09 panel pigtail | no ECO; open |
| E-26 | restated as a 6 mm tactile switch with a 12 mm coloured cap on an extender; panel openings 13.0 mm on a 14 mm pitch | correction, no ECO |
| E-27 | unchanged, and met in the design since 2 September 2026: the bicolour phase driver is written, alternating on the FreeRTOS tick at about 250 Hz against E-27’s “above 100 Hz”. No unit exists, so no light has ever been driven and T11 has not been run | ECO-EEG-001 |
| E-28 | restated as TP1–TP18 plus a 1×6 UART debug header at J26; the 2×5 1.27 mm JTAG/SWD header is withdrawn and every “E-28 deviation” note is deleted | correction, no ECO |
| E-29 | new: headphone output not to exceed 100 dB SPL at any commanded level on an artificial ear, with the codec volume register clamped at the calibrated value. Calculated full-scale output is about 110 dB SPL, which is why the requirement is needed. Added as a type test and to RISK-EEG-011 | new requirement, no ECO |
| F-06 | ring relaxed from three minutes to 90 seconds plus unlimited microSD backfill | ECO-EEG-025 |
| S-02 | unchanged, and met in the design since 2 September 2026 at a calculated 36.8 µA against 50 µA, R1–R16 having gone to 68 kΩ. Not signed off: RISK-EEG-011 SR-01 is closed in the design and the safety review has not started | ECO-EEG-024, implemented |
| S-04 | unchanged, and not met: no NTC net and no thermistor way exists | no ECO; open hardware item |
| S-09 | new: lithium shipping under UN3481 and PI967 | ECO-EEG-017 |
| A-03 | headband withdrawn as a kit item and rewritten to cover the HM-06 chin strap and the HM-03 occipital yoke, both on the frame | correction, no ECO |
| M-01 | POD-P1 enlarged to 163 × 143 × 58 mm external; the Phase 2 helmet-shell figures are sized for a carrier that no longer exists | ECO-EEG-018, first finding |
| M-02 | LED opening withdrawn: all eight lights are in the helmet and the pod carries no indicator | correction, no ECO |

### 2.3 Foreseen and not raised

Two things are known, are not defects today, and would become ECOs if a measurement goes the wrong way. They are listed here so that they are not discovered as surprises.

- The DevKit 3V3 rail. The carrier draws a calculated 288 mA worst case from the DevKit’s on-board regulator, which dissipates about 0.5 W in a small package inside a closed pod. Phase 1 measures it at TST-EEG-004 T3 and reports the regulator case temperature. If it exceeds 85 °C, a 3.3 V regulator on the carrier fed from V5V is an ECO against Rev C of the board. It is not solved today; it is measured.

- Conformal coating. The board is not coated for Phases 1 and 2, and that is a decision rather than a deferral: the board lives inside a gasketed enclosure, and coating a board with thirty connectors and a socketed DevKit costs more in masking than it buys. It is revisited before Phase 3 if a unit returns with corrosion.

## 2A. Corrections made by the third cross-document audit

A third pass, run after the second audit’s corrections were applied, looked for one thing only: a number, a citation or a claim in one document that a different document contradicts. It found thirteen. None of them is a design change – no schematic, netlist, layout, firmware or mechanical part moves – so none takes an ECO number. They are corrections, and they are listed here so that anyone holding a printed copy can see what changed and why.

| # | What was wrong | Where | Now |
| --- | --- | --- | --- |
| C-01 | SIM-EEG-018 graded the DRC as a pass while its own evidence column read “25 violations”, and the headline said “1 known open item” | tools/simulate_production.py | the check now asserts only that the violations are of the two known kinds, and the 25 are raised as an open item. History, and true when this row was written: the headline was 169 passed, 0 failed, 5 open. Superseded on 2 September 2026: the DRC reports zero violations, so the exception has nothing left to except and that open item is gone. The headline stood at 171 passed, 0 failed, 5 open when that sentence was written and is 193 passed, 0 failed, 6 open at the run behind this release – section 1.1 states it and names the six, and this row is not the place to read it from, which is the second time this row has had to be corrected for exactly the reason it warns about. A total written into a table as prose is a total that stops moving when the run does not; that is the second thing this row now guards. The first stands: a check that graded a DRC as a pass while its own evidence column read 25 violations is what it was raised against |
| C-02 | Two different byte rates were both called “the stream”: 50.0 kB/s in SIM, 50.7 kB/s everywhere else | SIM-EEG-018 | both are reported, named: 50.0 kB/s of raw sample payload, 50.7 kB/s framed (1014 bytes of frame, 1015 after COBS, every 20 ms) |
| C-03 | The ring buffer was 118 s in seven documents, 126 s in two | RUL-EEG-021, RFQ-EEG-001, TST-EEG-004, DSN-EEG-003, AVL-EEG-017, ECO-EEG-016, IFU-EEG-014 | 126 s. The 118 s divided 6 MB decimal by the framed rate; the ring is 6 MiB = 6,291,456 bytes and it holds raw samples. 124 s if counted framed. RING_BYTES and the 90 s of F-06 are unchanged, so this is arithmetic recorded under ECO-EEG-025, not a new ECO |
| C-04 | Ten documents said “no clearance violation” while the same page reported two of them. The Rev B closure named six and corrected three | corrected at Rev B: PARTS-EEG-019, the PCB spec sheet, README_package_index.txt. Named at Rev B but left unqualified until now: TST-EEG-004, ASM-EEG-007, RFQ-EEG-001. Never in the Rev B scope and also unqualified: ICD-EEG-006, QP-EEG-010, REG-EEG-012, FW-EEG-001 | each of the ten was made to say beyond those two electrode-net via clearances, which is what the DRC report supported at the time. Withdrawn on 2 September 2026: those two vias are gone and the report records no clearance violation at all, so the qualifier is now itself wrong and the plain sentence is correct. What survives is the rule, not the wording – no document states a DRC result it has not read in the report. The Rev B row was written as closed before three of the six it named had been edited; the scope is restated here rather than closed a second time. No automatic check guards the wording yet, so review holds it until one is added to tools/simulate_production.py; a copy of the sentence found without the qualifier is a reintroduction and is to be raised against this row |
| C-05 | SIM reported 156 surface-mount placements against 153 everywhere else | SIM-EEG-018 | 153 placed parts, of which R89 is do-not-populate; the CPL has 156 rows because it also carries the three fiducials, which take copper and mask apertures but no paste |
| C-06 | The PCB spec sheet said TST-EEG-004 has 30 steps; TST-EEG-004 says thirty-one | kicad/EEG-CAR-01_RevB_PCB_spec_sheet.txt | 31 steps (T00, T0, T1 to T29) |
| C-07 | RUL-EEG-021 printed the fiducials at (8, 8), (142, 8) and (8, 122), and three documents called that a transcription error while the register still carried it | RUL-EEG-021, DSN-EEG-003, ASM-EEG-007, QP-EEG-010 | the register carries (12, 10), (144, 100) and (12, 120), matching design.py and the CPL, and the withdrawal is stated in the register itself |
| C-08 | RUL-EEG-021 listed the ECOs as 001–015, 017 and 018, omitting the nine issued since | RUL-EEG-021 | 001–015, 017, 018 and 019–027, with 028 and 029 withdrawn |
| C-09 | RUL-EEG-021 said the isolation keep-out is clear on “either layer” | RUL-EEG-021 | all four layers |
| C-10 | SIM graded E-27 a pass on the light current alone, while ECO-EEG-016 and TST-EEG-004 T11 both record E-27 as not met | SIM-EEG-018 | the current check is retitled current only, and E-27’s missing bicolour phase driver is an open item. Superseded 2 September 2026: the driver is written, so the open item is no longer the missing driver but E-27 has never been seen to light – no unit exists and T11 has not been run. The split between a current check and a colour check is what this row bought, and it stands |
| C-11 | SIM graded E-22 on a 150 mA board current that ICD-EEG-006 disputes at about 440 mA | SIM-EEG-018 | E-22 is checked at both figures – 20.0 h and 6.8 h, so it is met either way – and the unreconciled current is an open item, as RFQ-EEG-001 Rev E open item 14 |
| C-12 | The DRC report’s “CONDUCTORS BELOW THE 0.25 mm PREFERRED WIDTH: 19” contained entries at 0.28 and 0.56 mm, and disagreed with the “200” under MEASURED | tools/drc.py | the section is CONNECTIONS THE ROUTER HAD TO RELAX, split into narrower than 0.25 mm and full width with a reduced gap, and it states that it counts connections while MEASURED counts segments. The counts move with the route: 169 as the board now stands, 36 and 133, against 19 when this row was written |
| C-13 | Four citations pointed at things that do not exist: firmware/pinmap_EEG-CAR-01.h (twice), “DSN-EEG-003, Rev B as it then stood, section 5”, “ASM-EEG-007 stage 2” for the reflow profile, and “FW-EEG-001 section 8” for F-08 | ICD-EEG-006, SIM-EEG-018 | firmware/main/board_pins.h, Rev C section 5, ASM-EEG-007 section 2.5, FW-EEG-001 section 5.6. The reflow figures now match ASM-EEG-007 section 2.5 exactly: 45–90 s above liquidus, 235–245 °C at the U1 body, 245 °C maximum anywhere |

### 2A.1 What package v1 had and package v2 had dropped

A fourth pass compared every file in package/ against package_v2.4/ by CONTENT, not by name – twenty-eight artefacts, including the four workbooks and the zip that sit beside the package folder. Twenty-three were covered outright. Five were not, and all five were cases of package v2 carrying a file forward in name while losing what was inside it. None was noticed by any of the three earlier audits, because all three read package v2 against itself and never opened package v1.

| # | What v1 had that v2 had lost | Severity | Now |
| --- | --- | --- | --- |
| C-14 | The entire internal cost model. EEG_kit_BOM_INTERNAL_costed.xlsx held an Assumptions sheet, 436 price cells across four quantity breaks, and a subsystem roll-up. EEG_kit_BOM_INTERNAL_RevB_costed.xlsx was a copy of the bidders’ workbook with two renamed columns and zero numeric price cells – it even carried the bidders’ footer, “no prices are given in this file by design” | major | Rebuilt as EEG_kit_BOM_INTERNAL_RevC_costed.xlsx from tools/emit_costed_bom.py, on v2’s 47-line part list, with the Assumptions sheet, per-line prices at 2/10/25/50, a per-line price basis column, the subsystem roll-up, the programme build-up with spares and NRE, and a v1-versus-v2 comparison. The empty Rev B file is deleted |
| C-15 | The mirrored solder-side view. v1 shipped EEG-CAR-01_placement_bottom.pdf. v2’s assembly drawing had a “bottom side” sheet, but it used the same transform as the top sheet, so it was the top view with a different title and every left-right instruction on it was wrong | major | tools/drawings.py gained a mirror transform. ASM 2 of 2 is now drawn as seen from underneath and says so on its face. This matters because ASM-EEG-007 §2.7 solders 217 socket-strip, 4 JST, 12 switch and 3 DIN joints from the bottom, and QP-EEG-010 IQC-B10 inspects against the sheet |
| C-16 | The nine manufacturers’ replies, and with them the only commercial evidence either package holds: Makerfabs’ blank-sheet ODM quotation of US$20-30k for Phase 1 design plus two samples | minor | Restored as DSN-EEG-003 §8.1, with the arithmetic that makes it matter: a blank-sheet engagement is four to seven times the whole of Phase 1 at the two-unit break, which is why the instrument is module-on-carrier |
| C-17 | The RFQ cover email, RFQ-EEG-001 Rev C Annex A – the only text in the package written for a reader who has opened nothing yet | minor | Restored as RFQ-EEG-001 Rev E Annex A, updated: four layers not two, and routing supplied not requested. It stated the twenty-five open items in the letter rather than leaving them to be discovered after quoting; since 2 September 2026 what it must state instead is that the data is released for review and not for fabrication, and that the routing is unreviewed. The reference-document list moves to Annex B |
| C-18 | Every foam corner radius. v1 specified sheet R14, HELMET R40, HEADPHONES/ELECTRODES R10 and the small pockets R8. Both v2 DXFs drew every profile square, and no document stated a radius | minor | tools/mech_gen.py draws true arcs at v1’s radii. PKG-EEG-015 §2.3 orders a steel-rule die for Phase 3, and steel rule cannot be folded to a true 90 degree inside corner – a square profile either tears the foam or silently becomes whatever radius the die maker picks |

Three findings raised by the same pass were refuted on re-check and are recorded here so they are not raised a fourth time: the netlist report’s two USB route rules (the geometry survives in v2 rule 6 and the stub prohibition in RFQ F-12), the Ø104 mm circle in v1’s foam top layer (a helmet dome relief made unnecessary by the two-layer split), and the three cover-email variants of v1’s Annex A (carried by DSN-EEG-003 Annex B’s conditional wording).

The lesson is the one C-01 to C-13 already taught, applied to a different axis. Those thirteen were package v2 contradicting itself. These five are package v2 contradicting package v1 – a file kept by name and emptied of content. A name-by-name diff finds none of them; only opening both files does.

Two further corrections were made at the same time and are not numbered above because they close a finding rather than resolve a contradiction: FW-D15 is marked closed, because the header it complained about now reads RFQ-EEG-001 Rev E; and SIM-EEG-018’s manual-handling note now states the two-to-three-minute fitting time as IFU-EEG-014 does – once the participant has done it once, with five minutes allowed for a first-timer.

## 2B. How the board was closed, 2 September 2026

Class: not a change to the instrument. No schematic, netlist, BOM, mechanical part or firmware line moves in this section: parts moved on the board and three faults in the programme’s own tools were corrected. It takes no ECO number. The placement comments in tools/design.py label these changes ECO-EEG-028, which is a number this register withdrew under section 2 and does not reuse; the record is here, and the next free ECO number is still ECO-EEG-030.

Was: the route ECO-EEG-018 was first written against – 122 of 145 nets connected, 23 unclosed, 7 of those with no copper at all, 25 DRC violations. Now: 145 of 145 connected, 0 unclosed, 0 without copper, 0 violations. Four things did it, and the order matters, because each one exposed the next.

- Placement, before routing. J4 moved 14 mm south, from y = 36 to y = 50, so that its analogue pins face the protection rows that feed them instead of fanning diagonally through one window and sealing each other in. R92 and R93, the separable AVDD2/AVSS2 rail links, moved off x = 54 – where the rail trunk ran between them and J29, which is why AVSS2 shipped with no copper at all – to the mid-latitude of the pins they serve, so no rail net runs the length of the socket. J22 moved from the bottom edge to directly under the protection ladder, which shortened two electrode runs from about 50 mm to under 9 mm and freed the ladder’s own escape lane. Smaller moves went with them: C80/C81, J9, J26, R94/R95, TP10/TP11, and one millimetre of extra pitch between the clamp and capacitor columns for a rail lane.

- The isolation keep-out was being measured from centrelines. The router blanked the cells of the strip but ignored the width of the copper it was carrying, so a 0.56 mm track whose centreline stayed outside the keep-out could still overhang it. The blank is now inflated per layer by the carried copper radius, so copper edges stay out of the strip. This is the mechanism behind the inner-plane copper that earlier revisions found inside it.

- The pre-placed reference vias were stamped on two layers of four. _place_ground_via stamped each one on the top layer and the first inner layer only, while Router.rebuild() stamps all four, so 69 vias were invisible to the bottom-layer router: it routed over them, and at AGND_REF pad U1.3 it ran a track through one and disconnected it. Stamping every reference via on all four layers is a one-line fix and it closed most of what was left open.

- Two targeted repairs. An AGND_REF stub outward from U1.3 with a through-via onto the planes, which reports itself skipped if the pad is already joined; and the SPARE2 corridor, R16.2 to D16.3, freed by ripping AVDD, AVSS, AGND_REF, EMGIN3 and EOGIN1 where they crossed it, routing SPARE2 first at its full electrode class and then re-joining all five ripped nets at their own full class. Rotating R16 was measured and rejected: pad 2 already faces D16, and turning it would move the pad away and push EOGIN2 across the lane SPARE2 needs.

What it cost, and it is not nothing. Connections the router had to relax rose from 19 to 169 – 36 with a conductor narrower than the 0.25 mm preferred width, 133 at full width with a reduced gap, every one at or above the 0.20 mm minimum conductor and gap, so none of them is a violation. Track segments went from 3 370 to 3 745 and through vias from 498 to 552. A board that closes at minimum geometry is not the same board as one that closes at preferred geometry, even when every rule passes. The 169 are listed in the DRC report by net and by pad pair, and they are the first thing RFQ-EEG-002A’s reviewer is asked to read.

Verified: kicad/EEG-CAR-01_RevB_DRC_report.txt – violations 0, 145 of 145 nets fully connected with none left without copper, minimum clearance 0.260 mm on F.Cu, 0.275 mm on B.Cu and 0.285 mm on the planes against a 0.200 mm rule, narrowest track 0.200 mm, smallest plated hole 0.300 mm, zero zone crossings, exactly one AGND_REF-to-DGND bridge and one HARN_SHIELD-to-DGND bridge, no duplicate segments and no duplicate via positions, and one continuous plane island per net on both inner layers.

### 2C Two findings of 2 September 2026, after the gate was met

Both came out of auditing the package against a question it had not been asked – whether anything attaches to the wrist – and neither is a routing defect.

The pad model. tools/fplib.py:171 gives every through-hole header pin but pin 1 the shape oval, which at equal sides is a circle; that is the KiCad convention and what kicad_write.py emitted. pours.py modelled every non-circular pad as a box and gerber.py apertured it R,1.6X1.6, so the connectivity check, the clearance check and the copper that would have been fabricated all carried a 1.6 mm square where the design has a 1.6 mm round pad. The corners are 27 % of the pad area and the router used them: 145 of 145 nets closed on the square, 62 of 145 on the round pad. Fixed – gerber.py now emits C,w for a square oval and pours.py models a circle, so the fabricated copper, the KiCad file and the check agree. 184 stranded track ends were given entry stubs to the pad centre by the new tools/pad_snap.py, added rather than substituted because moving an end rotates its segment and the first attempt put MISO 0.158 mm from a via of START. Re-checked: 145 of 145, VIOLATIONS: 0 on the round pads. Segments 3 560 -> 3 744, vias unchanged. That 3 744 is the figure at that repair and not the census of the released board, which is 3 745 segments and 552 vias – the bias-lead fix below added one of each. The routing-result table in ECO-EEG-018 and kicad/EEG-CAR-01_RevB_DRC_report.txt carry the current figures.

The bias-lead protection topology. As found: design.py wired all sixteen protection networks as src(patient) -[Rn]- dst(module) with Dn.3 and Cn.1 on dst, and then overrode channel 11 to put D11.3 and C11.1 on BIAS_EL, the patient side. A shorted D11 half therefore reached the Fpz electrode with no series resistor in the path at all (183.6 uA on this programme’s bound B, against S-02’s 50 uA), and a shorted C11 bonded that electrode to AGND_REF – the earth connection IFU-EEG-014 tells the participant does not exist. Rev B recorded it as not fixed, and not fixable by layout – RISK-EEG-011 SF-1a, SF-6a and SR-12 – with the sheet-2 schematic note that claimed every conductor reaching a person passes through one resistor and one clamp corrected against it.

Fixed on 2 September 2026, and it is a netlist change, not a layout one. The channel-11 exception in design.py is deleted: the row is generated by the same loop as the other fifteen, so D11.3 and C11.1 sit on BIASOUT, behind the series resistor, and the Fpz electrode is no longer reachable through a shorted clamp half with nothing in the path. SF-1a collapses into the ordinary SF-1, and no part is added and no value changes: an exception is removed. ECO-EEG-024 alone would not have closed this – D11 had no resistor in its fault path at all, so a larger resistor elsewhere changed nothing, which is why the two had to be taken together. Getting four re-labelled pads back into a closed board took a repair rather than a re-route: tools/repair_nets.py runs a full DRC afterwards and refuses to write unless the board comes out at least as good as it went in, and it refused twice. The board is 145 of 145 nets, 0 violations, 3 745 segments and 552 vias – one segment and one via more than before the fix. It is still for the electrical safety review that has not happened: RISK-EEG-011 owns SF-1, SF-6 and SR-12, and a fault path removed in the netlist is not a fault path reviewed.

Impact: the fabrication-release gate of section 3 is met, and the data in kicad/ is released for review under RFQ-EEG-002A; the routing-result table in ECO-EEG-018 above and section 3.4 of DSN-EEG-003 are restated from the report. Nothing else moves. Nothing here has been manufactured or measured, and no human layout engineer has reviewed the routing.

This is the record of 2 September 2026 and it stands as a record. Both of its last two statements are superseded by ECO-EEG-030 of 4 September 2026: a human layout engineer has read the routing, and the data in kicad/ for Rev B is withdrawn from fabrication rather than released for review. Section 3 is the current gate. Rev B closed this section by naming two requirements that were not met and were untouched by the routing work: S-02 at a calculated 53.2 µA, and S-04’s charge thermistor. Corrected 2 September 2026: S-02 is met in the design at 36.8 µA under ECO-EEG-024, and unsigned. S-04 is unchanged and still not met – there is no NTC net in design.py and no thermistor way on J12 or J13, so it does not exist and cannot be tested.

### 2D The firmware and the test tool, 2 September 2026

Class: not a change to the instrument, and it takes no ECO number. No requirement, schematic, netlist, BOM, mechanical part or pin assignment moves in this section. What moved is firmware source that did not do what FW-EEG-001 Rev D already required of it, and browser source that did not either. FW-EEG-001 owns the firmware defect register and the FW-Dnn numbers; this section records what happened so that the register is not the only place it is written down. The next free ECO number is unchanged at ECO-EEG-030.

The firmware could not have been built. firmware/main/drivers.h did not exist. main.c called seven functions that drivers.c defines – sd_free_mb, unit_serial_into, atecc_serial_into, battery_percent, envelope_onset_after, codec_play_tone_at and the drv_* group – with no declaration in scope. In C99 and later an implicit declaration is an error rather than a warning, and ESP-IDF compiles with -Werror=implicit-function-declaration, so the build stops. Where a toolchain does accept it, the failure is quieter and worse: an undeclared function is assumed to return int, so sd_free_mb()’s uint32_t and the pointer returns are truncated or misread at run time with no diagnostic. Now: drivers.h declares the API and is included by both main.c and drivers.c, so the compiler checks the definitions against the prototypes the callers use. main.c also now includes board_pins.h and esp_attr.h, which it had been using without including.

FW-D14: every command the browser sent started a recording session. A host command is a full frame – the ten-byte header of FW-EEG-001 section 5.1, then the opcode, then the opcode’s own payload. rx_task() handed handle_command() the frame from byte 0, so what the firmware read as the opcode was header byte 0, the protocol version. PROTO_VERSION is 1 and CMD_START_SESSION is 0x01, so IDENTIFY, LOOPBACK, CLOCK_XCHG and the rest all dispatched as START_SESSION, and the one command that appeared to work did so by that coincidence and not because anything was right. Now: rx_poll() strips the ten-byte header and validates the version and the frame type before dispatch, and the body of rx_task() is factored into rx_poll() so a host harness can drive the real decode path rather than a copy of it. FW-D14’s status in the defect register belongs to FW-EEG-001.

The acknowledgement did not have the shape the specification gives it. FW-EEG-001 section 6.2 fixes the CMD_ACK payload as opcode echoed, one reserved zero byte, status, result length, then the result. Every ack went out as two bytes with the status at offset 1 and no length, and IDENTIFY, LOOPBACK and CLOCK_XCHG skipped the shape altogether and returned their result at offset 0. That is not tidiness: with no opcode echoed on three replies, a host cannot tell which command an acknowledgement answers, and a queue that resolves the next pending command with whatever arrives will mis-attribute a reply under any load at all. Now: the firmware emits the section 6.2 shape, with defined statuses for ok, unknown opcode, the VBUS charge interlock of S-01 and not-implemented; and the browser tool’s transport.js and its simulated device were corrected to read and write the same shape.

Two capability bits said something they did not mean. CMD_IDENTIFY never set CAP_CODEC under any condition, although TOOL-EEG-022 section 2.3 defines bit 3 as codec initialised and the tool displays it, so a working codec always read as absent. CAP_ATECC was set only when the ATECC608B’s configuration zone was locked, conflating the chip is fitted with the unit is provisioned – different questions, asked at different stations, and the second is what T6 turns on. Now: CAP_CODEC is set when the codec is ready, CAP_ATECC when the chip answers, and CAP_PROVISIONED when the configuration zone is locked.

Why none of this was caught, which is the part worth keeping. The browser tool’s own test suites passed throughout, and they still pass. They could not have failed. The host and the simulated device it is tested against are both JavaScript, written together from the same reading of the specification, and they shared the same misunderstanding: the JS device answered an acknowledgement of { opcode, 0xFF }, the JS host expected exactly that, they agreed with each other on every run, and they agreed with neither FW-EEG-001 section 6.2 nor main.c. There were three acknowledgement shapes in the package – the specification’s, the firmware’s and the JavaScript pair’s – and the only pair ever compared was the one that could not disagree. The header defect was invisible for the same reason: the JS device read the header the JS host wrote, so nothing in that suite ever asked where the opcode was. It is not a weak test; it is a test of the wrong thing, and no quantity of it would have found FW-D14.

Now: webtest/tests/interop/ removes the shared author. It compiles the shipped firmware/main/main.c against small ESP-IDF stubs – on a copy, with one _Static_assert on a TinyUSB descriptor length neutralised because the stubs cannot expand that macro, and the shipped source is not modified – wires its USB endpoints to a pipe, and drives it with the browser tool’s own protocol.js and transport.js under Node. The C firmware and the JavaScript host are therefore checked against each other rather than each against its own reflection. It runs 57 checks, all of which pass and none of which fail – 32 when this section was written, and the count moved as checks were added, most recently for the provisioning opcodes of section 2E – and is invoked as sh webtest/tests/interop/run.sh. It needs a C compiler and Node, and neither ESP-IDF nor hardware. That total is the interop harness’s own; the production-simulation total is the one in section 1.1, and the two are not to be quoted for each other.

Impact: firmware source, the browser tool’s transport and simulated device, and one new test directory. Nothing here has been run on hardware – no unit exists, the interop harness runs the firmware’s dispatch path on a host compiler against stubs, and what it proves is that the two implementations agree with the specification and with each other, not that either works on the instrument. Rev B wrote “nothing here has been built or run on hardware”; the “built” half became false later the same day and is corrected in section 2E, which records the first real ESP-IDF build. The “run on hardware” half stands. Three documents carry the consequences and not this one: FW-EEG-001 owns the FW-Dnn register, the section 5 wire format and the section 6.3 opcode table that places CMD_IDENTIFY and CMD_LOOPBACK; TOOL-EEG-022 owns the tool’s own file list and its account of what the tool is tested against; TST-EEG-004 owns the T-numbers, and the interop harness has none and is not a production test step – TST-EEG-004 requires it re-run for every firmware image as a control on the tool, which is a different thing from a step a unit passes.

### 2E The firmware was built, and then it was run, 2 September 2026

Class: not a change to the instrument, and it takes no ECO number. No requirement, schematic, netlist, BOM, mechanical part or pin assignment moves here. What moved is the firmware project’s own build configuration and one host script. FW-EEG-001 owns the firmware, the FW-Dnn numbers and the opcode table; this section records what happened, because until 2 September 2026 no document in this package could say that the firmware had ever been through a compiler, and several said the opposite as a matter of record. The next free ECO number is unchanged at ECO-EEG-030.

#### The build

ESP-IDF v5.2.5, target esp32s3. firmware/release/ holds the four images ASM-EEG-007 section 6.1 and FW-EEG-001 section 9 tell the operator to flash – bootloader.bin at 0x0, partition-table.bin at 0x8000, ota_data_initial.bin at 0xF000 and eeg_field_kit.bin at 0x20000 – with manifest.json recording the SHA-256 of each, and the production simulation verifies every digest on every run. The application image is 405,360 bytes as flashed, and esp_idf_size reports 405,245 bytes of linked image in firmware/release/size.json. The build is the Phase 1 configuration, sdkconfig.defaults + sdkconfig.phase1: secure boot, flash encryption and anti-rollback off, so a prototype burns no eFuses and stays re-flashable (RUL-EEG-021 section B).

That closes a hole the package had carried since v1: the shop had a flashing procedure and nothing to run it on, while the same documents forbid the shop building the firmware itself.

#### The five defects the build found, none of which any check in this package could have found

The interop harness of section 2D compiles main.c against stubs, and a stub agrees with anything. A real toolchain does not.

- esp_driver_i2c does not exist at ESP-IDF 5.2. The I²C driver was split into its own component at 5.3 and main/idf_component.yml permits 5.2, so the project required a component by name at a version where it is absent, and configuration stopped dead. The driver component provides I²C at both versions.

- sdkconfig.phase1 set CONFIG_BOOTLOADER_APP_ANTI_ROLLBACK twice – =n, with a comment explaining why a Phase 1 prototype has none, and =y eleven lines later in a block added to close FW-D11. The later duplicate silently wins, so the Phase 1 intent was being undone by a line that read as an addition. It also fails the build outright: ESP-IDF refuses anti-rollback on a partition table carrying a factory partition, and partitions.csv has one. That is how it was found – the first real build stopped there.

- CONFIG_TINYUSB_VENDOR_ENABLED does not exist. esp_tinyusb 1.4.x calls it TINYUSB_VENDOR_COUNT, and ESP-IDF warns on an unrecognised key rather than failing, so the setting looked present while CFG_TUD_VENDOR stayed 0, tusb.h declared none of the tud_vendor_* functions, and the WebUSB interface this whole instrument is built around would not have existed. It surfaced as an implicit declaration of tud_vendor_mounted the first time the firmware met a real compiler.

- Three descriptor callbacks were defined twice. main.c defined tud_descriptor_device_cb, _configuration_cb and _string_cb and passed the same descriptors through tinyusb_config_t; esp_tinyusb defines those three itself, so the link failed with three multiple-definition errors. With CONFIG_TINYUSB_DESC_CUSTOM the descriptors are data and the component owns the callbacks. tud_descriptor_bos_cb stays, because the component defines no BOS callback and without BOS there is no WebUSB.

- firmware/main/drivers.h did not exist, so main.c called seven functions with no declaration in scope. That one was found earlier the same day and is recorded in section 2D above; it is counted here because it is the same class of defect and the same compiler would have caught it.

One more thing the build measured rather than fixed. Static IRAM links with one byte free of 16,384. That is a cliff, not a pass: the next function anyone marks IRAM_ATTR fails the link with an error naming a section rather than a cause, and this design has more interrupt work coming. Two ISRs this firmware does not use – there is no SPI slave in the design and nothing uses gptimer – were taken out of IRAM and the figure did not move by a single byte, so that is not the cause and is not to be recorded as the fix. It is open item 1 of the simulation run, and settling it means reading the linker map against hardware.

#### The QEMU run, and what it does not prove

firmware/release/qemu_boot.log is one full boot cycle from qemu-system-xtensa 9.0.0, -M esp32s3, against this build with an emulation overlay – sdkconfig.qemu, which sets CONFIG_SPIRAM_IGNORE_NOTFOUND because QEMU’s esp32s3 has no octal PSRAM, and which is an emulation overlay and not a build configuration for any unit. What executed, none of which ever had: the second-stage bootloader read this partition table, all nine entries, and loaded the app from the factory slot at 0x20000; app_main() reached its own banner; the microSD mount failed, as it must with no card, and the firmware continued with the warning it was written to give; the ES8388 failed and degraded the same way; and the 6 MiB ring would not allocate, where FW-D13’s NULL guard caught it and named the cause, CONFIG_SPIRAM_MODE_OCT and the -N16R8 part. Three bring-up paths exercised, three behaved as designed.

It has never run on hardware, and QEMU carries none of the peripherals this firmware talks to – no octal PSRAM, no microSD, no ES8388, no ADS1299. No register value, no daisy-chain order and no SPI timing has met silicon. The run proves the part that needs none of them. About the instrument it proves nothing.

And it found a failure mode, which is what running things is for. Two subsystems degrade and one aborts: a missing card and a missing codec are warnings, a missing ring calls abort(), and the unit reboots into the same abort for ever. On a bench that is right – without the ring there is no recording. In a participant’s home it is the worst available behaviour: a boot loop is indistinguishable from a flat battery or a broken cable, the device never enumerates, and the browser tool that exists precisely to answer “what is wrong with this unit” can never reach it. It is raised against F-06 and TOOL-EEG-022 and it is not closed here.

#### The provisioning opcode collision, and the guard that was missing

firmware/tools/provision.py and main.c both allocated 0x4A on 2 September 2026 – the host for the ATECC configuration-zone write, the firmware for CMD_READ_CALIBRATION – because each picked the next free opcode without the other. Moving the host’s config write to 0x4B then landed on an opcode the same file was already using, so for a short time both CMD_ATECC_WRITE_CONFIG and CMD_READ_CALIBRATION stood at 0x4B: a second collision, made while fixing the first, and this one inside a single file. Provisioning step 7b’s read-back would have sent a three-byte request that the firmware parses as a 65-byte configuration write and answers bad-length, so TST-EEG-004 T6’s acceptance read-back would have failed on every unit.

Both collisions are closed, and main.c is the authority: 0x4A reads calibration, 0x4B writes the configuration zone, which is what the firmware implements and what provision.py now sends. The guard matters more than the fix. provision_selftest.py now asserts that every CMD_ constant in that file is unique and matches main.c – the check whose absence let one opcode be allocated twice in a single day – and two of the interop harness’s 57 checks pin the meanings from the other end, refusing a short config write with bad-length and confirming that 0x4A still answers as the calibration reader. An opcode table belongs in one place; until it is in one place, a test has to hold it.

Impact: the firmware project’s component list and build configuration, one host script and its self-test, and four release images with a manifest. Nothing here changes the instrument, and nothing here is a bring-up. Three documents carry the consequences and not this one: FW-EEG-001 owns the FW-Dnn register, the section 6.3 opcode table and the ESP-IDF version it pins; TOOL-EEG-022 owns what the tool is tested against; TST-EEG-004 owns T6 and the T-numbers. Any document still saying that the firmware has never been compiled is wrong as of 2 September 2026, and each of those documents corrects that for itself – this register records the fact, it does not amend other documents from here.

## 3. Change procedure from here

- Raise. Anyone may raise a change. It is written as one paragraph: what is wrong, how it was found, and what it stops someone doing.

- Classify. Blocking – something cannot be built, tested or used safely. Major – a requirement is not met, or a document says two different things. Minor – everything else.

- Assess. State the impact on: the board, the firmware, the mechanical parts, the BOM, the test specification, the safety case, and the cost of any units already built.

- Change the source, not the output. If the change touches the carrier, edit tools/design.py and run tools/emit_all.py. If it touches a printed part, edit tools/mech_gen.py. Never edit a generated file.

- Verify. Re-run tools/emit_all.py and tools/simulate_production.py and attach both reports to the ECO. Two gates apply, and they are not the same gate.

- The gate for releasing fabrication data is that the DRC reports zero violations, that every net is one connected copper island, and that both inner planes remain continuous under the analogue zone. Rev B met all three on 2 September 2026, and that is recorded in section 2B. It is not enough, and this paragraph is rewritten rather than left standing. An external layout engineer read the Rev B board between 3 and 5 September 2026, declined the review and advised a redesign; his seven findings are in ECO-EEG-030 and six of them are rules the programme’s own DRC was never given, so a report at zero violations was a report against an incomplete rule set. Rev B is withdrawn from fabrication. No board is ordered from it and its data stays in the tree as history.

- Nothing in this package is released for fabrication. Rev C is defined by ECO-EEG-030, is unrouted, and is released to an external layout contractor under RFQ-EEG-002A as re-scoped by that ECO. The gate for releasing Rev C fabrication data is the three conditions above measured on the geometry the contractor returns, under the rule set of ECO-EEG-032 and LAY-EEG-034 rather than the Rev B rule set, plus the placement review of RFQ-EEG-002A closed before routing was confirmed. Until a contractor’s board has been returned and graded, there is no candidate for the gate at all.

- The gate for an ECO inside this release is narrower, because the release is a review package rather than a fabrication package: the change must add no violation, must not increase the count of nets that are not one island, must not put copper in the isolation keep-out on any layer, and every violation it closes must be named in the ECO. Now that the report is at zero, a change that adds one is a regression and is reverted rather than argued. A change that adds a CHECK is not in that class: ECO-EEG-032 adds six geometry checks that Rev B was never measured against, and the counts they raise on the Rev B routing are evidence that the checks bite, not regressions to be reverted. The simulator must stay at zero failures; its open items are counted and named. The DRC item is gone, because the report records no violations. The six that stand are the static IRAM pool with one byte free, the carried-over v1 HM-01 mesh being two disconnected bodies, the unreviewed routing, E-27 never having been seen to light, the unreconciled board current, and SR-01 being closed in the design and not signed off – 193 passed, 0 failed, 6 open, which is what the last run of 2 September 2026 printed and what section 1.1 states. E-11’s low-pass half and E-27’s missing driver were on this list at the previous run and are gone from it because both were closed, not because either was regraded. A change that raises the failure count is a regression. A change that moves the passed count moves it because the check set moved, and the ECO says which checks it added or withdrew.

- Record. Add the ECO to section 2 with the same fields as the entries above and take the next free number in the one EEG-nnn pool that section 1.3A now rules, which is ECO-EEG-034. Never ECO-EEG-016, which is this document, and never ECO-EEG-028 or ECO-EEG-029, which are withdrawn under section 2 and are not reused. Bump the revision of every document the change touches, and update sections 1.1, 1.5 and 2.1.

- Release. A release is the whole package or nothing. Regenerate the checksums in kicad/gerber/README_layer_map_and_checksums.txt and the package manifest.

A change that affects the safety case – anything touching the electrode path, the isolation, the battery, the charge interlock or the patient-current budget – additionally requires RISK-EEG-011 to be re-issued and the safety reviewer to be told, whether or not the review has already happened. ECO-EEG-023 and ECO-EEG-024 are both in that class and both are open, so the safety reviewer receives them as part of the pack rather than as changes already made.

A change that needs a test step cites a T-number from TST-EEG-004, which owns the step numbers. If the step does not exist, the ECO says so and raises it as an open item; it does not invent a number.

## 4. What is deliberately not under change control

The experiment design document, the participant information sheet and the pre-registration are the programme’s research documents and are controlled elsewhere. They are cited from this package but are not part of it. Where package v2 documents cited them as ETH-EEG-001, DIS-EEG-001 or KPL-EEG-001, those identifiers are withdrawn under section 1.3 and the documents are cited by title.