Table of Contents

# End-to-end production simulation

Document: SIM-EEG-018 Revision: A Date: 2026-09-02 Issued by: TI One Voice research programme (one.witysk.org), Brussels, Belgium Licence: CC BY-SA 4.0 Generated by: tools/simulate_production.py, from tools/design.py and the files actually present in the package

## What this is, and what it is not

This report is a dry run of the manufacturing route against the data package. At each station it asks what the operator needs, checks that the package contains it, and where the station has an acceptance limit it computes that limit from component values instead of quoting it. It is re-run by tools/simulate_production.py and must be clean before the package is released.

It cannot find what only a real build finds: solderability, fit, ergonomics, electromagnetic behaviour, or whether a participant can actually put the helmet on. Nothing in this package has been manufactured or measured.

Result: 193 checks passed, 0 failed, 7 known open items.

A failure means the package does not support the step. An open item means it does, and somebody still has to take a decision that is recorded where it belongs. The open items are:

| Station | Item | Detail |
| --- | --- | --- |
| 01 bare-board fabrication | static IRAM reports full, with one byte free | 1 byte(s) free of 16,384 on the linked image. If that is the real limit then the next function marked IRAM_ATTR fails the link with an error naming a section and not a cause, and this design has more interrupt work coming (E-13’s tone scheduler, E-12’s onset detector). Turning off the SPI-slave and gptimer ISRs – neither of which this firmware uses – did not move the figure by a single byte, so it is not those. Either the pool is genuinely full, or esp_idf_size is reporting against a fixed 16 KB window that is not the limit on an ESP32-S3 with octal SPIRAM and XIP. Reading the linker map against hardware settles it; guessing at sdkconfig from here does not |
| 01 bare-board fabrication | the v1 HM-01 mesh is two disconnected bodies | measured with trimesh: it is watertight, which MANIFEST.json records, and it is not connected, which nothing recorded. It is a carried-over form study rather than a released printable part – the parametric channel runs HM-01P-A/B/C were measured from it – but a frame that is two shells is one more reason the redraw of PARTS-EEG-019 OA-1 has to happen |
| 01 bare-board fabrication | there is no fabrication data for the current board | Rev B is withdrawn from fabrication (ECO-EEG-030) and Rev C is UNROUTED: its placement and routing are bought. The Gerbers, drill, DRC report and drawings this station checks are Rev B’s and are kept as history. A bare board cannot be ordered. What exists for Rev C is the layout input set in kicad/RevC_layout_inputs/ and the rule sheet LAY-EEG-034 |
| 01 bare-board fabrication | the placement of Rev C has not been reviewed | A human layout engineer HAS now read the Rev B routing: he read it between 3 and 5 September 2026, declined the paid review and advised a redesign, and Rev B is withdrawn (ECO-EEG-030). What is open is the next review, not that one. RFQ-EEG-002A is re-scoped to the review of the layout contractor’s PLACEMENT of Rev C, taken at the placement confirmation gate and before routing is confirmed, as a written findings list by net, pad pair and coordinate. There is no placement to review yet |
| 09 functional test TST-EEG-004 | E-27 has never been seen to light | the driver is written and the current budget is met, but no unit exists, so no light has been driven and TST-EEG-004 T11 – which reads the R/G ratio with a colorimeter – has not been run. The alternation also quantises to the FreeRTOS tick, about 250 Hz rather than exactly 240, which meets the ‘above 100 Hz’ E-27 is written against and is what T11 will actually measure |
| 09 functional test TST-EEG-004 | the two board-current figures cannot both be right | TST-EEG-004 T3 limits J13 to 150 mA while ICD-EEG-006 section 2.7 tallies about 440 mA. E-22 is met either way (20.0 h against 6.8 h), but the charger, the thermal budget and the T3 limit itself all rest on the number, and it is open item 14 of RFQ-EEG-001 Rev E. It is measured, not calculated, at T3 |
| 09 functional test TST-EEG-004 | SR-01 is closed in the design and not yet signed off | S-02 is now met at 36.8 uA against 50 uA, because ECO-EEG-024 raised R1-R16 to 68 kOhm. Applying the fix the analysis pointed to is not the same as having it approved: the electrical safety reviewer of RISK-EEG-011 section 7 owns SR-01 and that review has not started. |

## Station by station

### Station 00 purchase order and incoming goods

|  | Check | Detail |
| --- | --- | --- |
| pass | carrier BOM exists | /Users/nstephane/Dev/EEG-kit/kicad/EEG-CAR-01_RevC_BOM.csv |
| value | BOM lines | 79 |
| pass | every purchased line carries a manufacturer part number |  |
| pass | BOM quantity equals the placed part count | BOM 207, placed 207 |
| value | do-not-populate | R89 |
| pass | approved vendor list present |  |
| pass | incoming inspection defined (quality plan) |  |

### Station 01 bare-board fabrication

|  | Check | Detail |
| --- | --- | --- |
| pass | EEG-CAR-01-F_Cu.gbr present |  |
| pass | EEG-CAR-01-In1_Cu.gbr present |  |
| pass | EEG-CAR-01-In2_Cu.gbr present |  |
| pass | EEG-CAR-01-B_Cu.gbr present |  |
| pass | EEG-CAR-01-F_Mask.gbr present |  |
| pass | EEG-CAR-01-B_Mask.gbr present |  |
| pass | EEG-CAR-01-F_Silkscreen.gbr present |  |
| pass | EEG-CAR-01-B_Silkscreen.gbr present |  |
| pass | EEG-CAR-01-Edge_Cuts.gbr present |  |
| pass | EEG-CAR-01-PTH.drl present |  |
| pass | EEG-CAR-01-NPTH.drl present |  |
| pass | EEG-CAR-01-F_Cu.gbr is a complete Gerber (FS, MO, M02) |  |
| pass | EEG-CAR-01-F_Cu.gbr uses only apertures it defines |  |
| pass | EEG-CAR-01-F_Cu.gbr coordinates lie inside the board outline | x 2.90..148.00 y 1.90..128.50 |
| pass | EEG-CAR-01-In1_Cu.gbr is a complete Gerber (FS, MO, M02) |  |
| pass | EEG-CAR-01-In1_Cu.gbr uses only apertures it defines |  |
| pass | EEG-CAR-01-In1_Cu.gbr coordinates lie inside the board outline | x 0.40..149.60 y 0.40..129.60 |
| pass | EEG-CAR-01-In2_Cu.gbr is a complete Gerber (FS, MO, M02) |  |
| pass | EEG-CAR-01-In2_Cu.gbr uses only apertures it defines |  |
| pass | EEG-CAR-01-In2_Cu.gbr coordinates lie inside the board outline | x 0.40..149.60 y 0.40..129.60 |
| pass | EEG-CAR-01-B_Cu.gbr is a complete Gerber (FS, MO, M02) |  |
| pass | EEG-CAR-01-B_Cu.gbr uses only apertures it defines |  |
| pass | EEG-CAR-01-B_Cu.gbr coordinates lie inside the board outline | x 2.70..147.09 y 3.50..127.80 |
| pass | EEG-CAR-01-F_Mask.gbr is a complete Gerber (FS, MO, M02) |  |
| pass | EEG-CAR-01-F_Mask.gbr uses only apertures it defines |  |
| pass | EEG-CAR-01-F_Mask.gbr coordinates lie inside the board outline | x 5.00..145.94 y 3.50..126.70 |
| pass | EEG-CAR-01-B_Mask.gbr is a complete Gerber (FS, MO, M02) |  |
| pass | EEG-CAR-01-B_Mask.gbr uses only apertures it defines |  |
| pass | EEG-CAR-01-B_Mask.gbr coordinates lie inside the board outline | x 5.00..144.00 y 17.14..125.00 |
| pass | EEG-CAR-01-F_Silkscreen.gbr is a complete Gerber (FS, MO, M02) |  |
| pass | EEG-CAR-01-F_Silkscreen.gbr uses only apertures it defines |  |
| pass | EEG-CAR-01-F_Silkscreen.gbr coordinates lie inside the board outline | x 3.00..146.03 y 4.85..128.38 |
| pass | EEG-CAR-01-B_Silkscreen.gbr is a complete Gerber (FS, MO, M02) |  |
| pass | EEG-CAR-01-B_Silkscreen.gbr uses only apertures it defines |  |
| pass | EEG-CAR-01-B_Silkscreen.gbr coordinates lie inside the board outline | x 17.92..61.91 y 7.00..10.60 |
| pass | EEG-CAR-01-Edge_Cuts.gbr is a complete Gerber (FS, MO, M02) |  |
| pass | EEG-CAR-01-Edge_Cuts.gbr uses only apertures it defines |  |
| pass | EEG-CAR-01-Edge_Cuts.gbr coordinates lie inside the board outline | x 0.00..150.00 y 0.00..130.00 |
| pass | EEG-CAR-01-PTH.drl is a complete Excellon file (M48, M30) |  |
| pass | EEG-CAR-01-PTH.drl declares metric units |  |
| pass | EEG-CAR-01-PTH.drl defines every tool it uses | undeclared: [] |
| value | EEG-CAR-01-PTH.drl tools | 0.300 mm, 0.900 mm, 1.000 mm, 1.200 mm, 1.700 mm |
| pass | EEG-CAR-01-NPTH.drl is a complete Excellon file (M48, M30) |  |
| pass | EEG-CAR-01-NPTH.drl declares metric units |  |
| pass | EEG-CAR-01-NPTH.drl defines every tool it uses | undeclared: [] |
| value | EEG-CAR-01-NPTH.drl tools | 1.500 mm, 3.200 mm |
| pass | non-plated hole count matches the design | drill 10, design 10 |
| pass | no hole appears in both drill files |  |
| value | plated holes | 788 |
| pass | the pad the DRC measures is the pad the design declares | all pads agree |
| pass | every pin-1 legend mark is nearest pad 1 | all footprints agree |
| pass | IPC-D-356A netlist present |  |
| pass | every IPC-D-356A record is the same width | widths [74] |
| pass | every IPC-D-356A drill diameter is one the design has | 5 distinct, all matched |
| pass | solder paste layer supplied |  |
| pass | every stencil aperture is one ASM-EEG-007 section 2.3 specifies | 7 distinct apertures, all from the table |
| pass | every land has an aperture rule | all lands ruled |
| pass | every aperture is above the 0.66 area-ratio floor at 0.12 mm | worst 1.53 at (0.95, 0.6) mm |
| note | mechanical fit | cadquery is not installed, so the HM-04/HM-05B bayonet fit was not measured on this run |
| note | frame section | cadquery is not installed, so the HM-01P channel section was not measured on this run |
| pass | the firmware release images are in the package | four images and a manifest |
| pass | every image matches the SHA-256 in its manifest | 4 images |
| note | firmware release | Built against a real ESP-IDF for the first time on 2026-09-02. Getting here took five corrections to the project, each recorded in the file it was made in: a component that does not exist at this IDF version, a duplicated anti-rollback key that contradicted the Phase 1 intent and made the build refuse a table with a factory partition, a TinyUSB option name that does not exist so the vendor interface was silently absent, and three descriptor callbacks defined here AND in esp_tinyusb. NOT RUN ON HARDWARE: no board exists. This is a build, not a bring-up. Re-issued 2026-09-02 after FW-D17 and FW-D18: the contact-light driver now derives its three colours from two lead-off comparator thresholds rather than from LOFF_STATP and LOFF_STATN, because this board’s montage is single-ended and the N half carried no per-site information, which had made the red state unreachable; and CMD_LIGHTS modes 2, 3 and 4 are implemented rather than silently answered OK. |
| value | firmware image | 405,541 bytes; static IRAM 16,383 used, 1 free |
| open | static IRAM reports full, with one byte free | 1 byte(s) free of 16,384 on the linked image. If that is the real limit then the next function marked IRAM_ATTR fails the link with an error naming a section and not a cause, and this design has more interrupt work coming (E-13’s tone scheduler, E-12’s onset detector). Turning off the SPI-slave and gptimer ISRs – neither of which this firmware uses – did not move the figure by a single byte, so it is not those. Either the pool is genuinely full, or esp_idf_size is reporting against a fixed 16 KB window that is not the limit on an ESP32-S3 with octal SPIRAM and XIP. Reading the linker map against hardware settles it; guessing at sdkconfig from here does not |
| pass | every printed part is a single closed body | 29 meshes, all one body |
| open | the v1 HM-01 mesh is two disconnected bodies | measured with trimesh: it is watertight, which MANIFEST.json records, and it is not connected, which nothing recorded. It is a carried-over form study rather than a released printable part – the parametric channel runs HM-01P-A/B/C were measured from it – but a frame that is two shells is one more reason the redraw of PARTS-EEG-019 OA-1 has to happen |
| pass | cross-language interop harness supplied |  |
| pass | layer map and checksums supplied |  |
| pass | Rev B fabrication drawing retained as history |  |
| pass | the Rev B geometry has been regraded under the ECO-EEG-032 rule set |  |
| pass | the layout rule sheet is issued |  |
| pass | the 3D collision check has been run |  |
| open | there is no fabrication data for the current board | Rev B is withdrawn from fabrication (ECO-EEG-030) and Rev C is UNROUTED: its placement and routing are bought. The Gerbers, drill, DRC report and drawings this station checks are Rev B’s and are kept as history. A bare board cannot be ordered. What exists for Rev C is the layout input set in kicad/RevC_layout_inputs/ and the rule sheet LAY-EEG-034 |
| pass | the Rev C layout input set is assembled |  |
| pass | the Rev C schematic matches the design source | tools/sch_netlist.py reads the emitted .kicad_sch back and diffs it against design.py |
| pass | every footprint is audited against its part number | docs/footprint_audit_RevC.md |
| pass | Rev B DRC report retained as history |  |
| pass | the Rev B DRC reported no violations against ITS rule set | VIOLATIONS: 0. That rule set did not contain six of the seven findings of ECO-EEG-030; the same geometry regraded under them is in kicad/EEG-CAR-01_RevB_regraded_ECO-EEG-032.txt |
| pass | every routable net is one connected copper island | 145 of 145 connected, 0 unclosed, 0 without copper |
| value | nets fully connected | 145 of 145 routable (156 in the design) |
| value | connections routed at relaxed geometry | 169 |
| note | release state | the ECO-EEG-016 section 3 gate is MET – zero violations, every net one connected copper island. The fabrication data is RELEASED FOR REVIEW under RFQ-EEG-002A and is NOT released for fabrication: no human layout engineer has read this routing, and 169 of its connections close at the minimum conductor or the minimum gap rather than the preferred width |
| open | the placement of Rev C has not been reviewed | A human layout engineer HAS now read the Rev B routing: he read it between 3 and 5 September 2026, declined the paid review and advised a redesign, and Rev B is withdrawn (ECO-EEG-030). What is open is the next review, not that one. RFQ-EEG-002A is re-scoped to the review of the layout contractor’s PLACEMENT of Rev C, taken at the placement confirmation gate and before routing is confirmed, as a written findings list by net, pad pair and coordinate. There is no placement to review yet |

### Station 02 bare-board electrical test

|  | Check | Detail |
| --- | --- | --- |
| pass | IPC-D-356A netlist supplied |  |
| pass | every net appears in the test netlist |  |
| pass | netlist terminates with 999 |  |
| value | test points in the netlist | 1172 |

### Station 03 SMT assembly

|  | Check | Detail |
| --- | --- | --- |
| pass | SMT pick-and-place file supplied |  |
| pass | every surface-mount part is in the CPL | missing: [] |
| pass | CPL origin is the bottom-left corner, Y positive | x 12.0..145.0 y 7.0..126.7 |
| pass | every CPL row states the layer |  |
| value | SMT placements | 153 placed parts, of which R89 is do-not-populate; the CPL has 156 rows because it also carries the 3 fiducials, which take copper and mask apertures but no paste |
| value | distinct SMT packages | 8: C_0603_1608Metric, Fiducial_1mm_Mask3mm, L_0603_1608Metric, R_0603_1608Metric, R_1206_3216Metric, SOT-23, SOT-23-5, TSSOP-14_4.4x5mm_P0.65mm |
| pass | stencil (paste) layer supplied |  |
| pass | every surface-mount pad that takes paste has a stencil aperture | stencil 366, pads taking paste 366 |
| value | surface-mount pads deliberately without paste | 24 (TP1-TP18 probe pads and the three fiducials: copper and mask only, nothing is soldered to them) |
| note | reflow profile | SAC305: preheat 150-200 C over 90 s, 45-90 s above the 217 C liquidus, peak 235-245 C measured at the U1 body and 245 C maximum anywhere on the board, cooling <= 4 C/s – ASM-EEG-007 section 2.5 |
| pass | assembly work instructions present |  |

### Station 04 through-hole assembly

|  | Check | Detail |
| --- | --- | --- |
| pass | through-hole position file supplied |  |
| pass | every through-hole part is listed | missing: [] |
| value | through-hole parts | 33 |
| pass | R90 and R91 exist and are the only star points |  |
| pass | no component other than R90 bridges AGND_REF and DGND | bridging parts: [‘R90’] |
| pass | no component other than R91 bridges HARN_SHIELD and DGND | bridging parts: [‘R91’] |

### Station 05 module preparation and the MP-01 plate

|  | Check | Detail |
| --- | --- | --- |
| pass | interface control document present |  |
| pass | module plate MP-01 model present |  |
| pass | J1 exists on the carrier | ADS1299 module #1 digital |
| pass | J2 exists on the carrier | ADS1299 module #1 analogue signals |
| pass | J23 exists on the carrier | ADS1299 module #1 analogue rails |
| pass | J3 exists on the carrier | ADS1299 module #2 digital |
| pass | J4 exists on the carrier | ADS1299 module #2 analogue signals |
| pass | J29 exists on the carrier | ADS1299 module #2 analogue rails |
| pass | J6 exists on the carrier | ESP32-S3-DevKitC-1 row A |
| pass | J7 exists on the carrier | ESP32-S3-DevKitC-1 row B |
| pass | J8 exists on the carrier | ES8388 codec module |
| pass | J9 exists on the carrier | Codec microphone feeds |
| pass | J10 exists on the carrier | ADuM4160 USB isolator, device side |
| pass | J11 exists on the carrier | ATECC608B breakout |
| pass | J12 exists on the carrier | Charger + fuel gauge module |
| pass | J19 exists on the carrier | 74HC595 contact-light driver |
| pass | J20 exists on the carrier | microSD breakout, SDMMC 1-bit |
| pass | J21 exists on the carrier | Boom-mic preamp module (part not settled, E-14) |
| pass | J25 exists on the carrier | Buck-boost module, TPS63020 class |
| pass | J28 exists on the carrier | Room microphone module |
| pass | DevKitC-1 header row spacing is 22.86 mm | 22.86 mm |
| pass | J7 position 11 is left unconnected (PSRAM / strapping) | NC_GPIO37 |
| pass | J7 position 12 is left unconnected (PSRAM / strapping) | NC_GPIO36 |
| pass | J7 position 13 is left unconnected (PSRAM / strapping) | NC_GPIO35 |
| pass | J7 position 15 is left unconnected (PSRAM / strapping) | NC_GPIO45 |

### Station 06 harness and cable assembly

|  | Check | Detail |
| --- | --- | --- |
| pass | harness specification present |  |
| pass | electrode harness carries 8 scalp + 2 reference + bias + screen = 12 ways | E_Fz, E_Cz, E_Pz, E_C3, E_C4, E_T7, E_T8, E_F7, REF_L, REF_R, BIAS_EL, HARN_SHIELD |
| pass | light harness carries 8 lines + common + return = 10 ways | LED1, LED2, LED3, LED4, LED5, LED6, LED7, LED8, LED_V, LED_GND |
| pass | no digital signal shares the electrode harness |  |

### Station 07 mechanical and enclosure

|  | Check | Detail |
| --- | --- | --- |
| pass | HM-04_electrode_assembly_body.stl present |  |
| pass | HM-04_electrode_assembly_body.step present (a mesh cannot be dimensioned) |  |
| pass | HM-08_battery_hatch.stl present |  |
| pass | HM-08_battery_hatch.step present (a mesh cannot be dimensioned) |  |
| pass | HM-09_service_key.stl present |  |
| pass | HM-09_service_key.step present (a mesh cannot be dimensioned) |  |
| pass | HM-02_brow_pad.stl present |  |
| pass | HM-02_brow_pad.step present (a mesh cannot be dimensioned) |  |
| pass | FIT-01_fit_test_coupon.stl present |  |
| pass | FIT-01_fit_test_coupon.step present (a mesh cannot be dimensioned) |  |
| pass | POD-P1_prototype_enclosure_base.stl present |  |
| pass | POD-P1_prototype_enclosure_base.step present (a mesh cannot be dimensioned) |  |
| pass | POD-P1_prototype_enclosure_lid.stl present |  |
| pass | POD-P1_prototype_enclosure_lid.step present (a mesh cannot be dimensioned) |  |
| pass | MP-01_module_plate.stl present |  |
| pass | MP-01_module_plate.step present (a mesh cannot be dimensioned) |  |
| pass | the CASE-00 Rev C foam cut files are supplied, all seven layers | 7 layer file(s) in mech/ |
| pass | no superseded Rev B foam file is still shipped | PKG-EEG-015 2.4: the Rev B two-layer insert cannot pack this kit |
| pass | 2D drawings supplied for the printed parts |  |
| pass | rulings register present |  |
| pass | part identifier register present |  |
| pass | carrier fits the POD-P1 internal envelope (158 x 138 mm) | board 150.0 x 130.0 mm |
| pass | the electronics stack fits the 55.5 mm internal depth | 49.1 mm: floor 2.5 + boss 6.0 + carrier 1.6 + standoff 18.0 + plate 3.0 + modules 18.0 |
| value | mounting hole pattern | M3 at (5,5) (145,5) (5,125) (145,125) – 140 x 120 mm centres, shared by the carrier, MP-01 and the POD-P1 bosses |

### Station 08 firmware build, flash and provisioning

|  | Check | Detail |
| --- | --- | --- |
| pass | firmware/main/main.c present |  |
| pass | firmware/CMakeLists.txt present |  |
| pass | firmware/main/CMakeLists.txt present |  |
| pass | firmware/sdkconfig.defaults present |  |
| pass | firmware/partitions.csv present |  |
| pass | firmware/README.md present |  |
| pass | provisioning script present |  |
| pass | host verification tool present |  |
| pass | generated pin map header present |  |
| pass | no signal is assigned to GPIO35 (octal PSRAM) |  |
| pass | no signal is assigned to GPIO36 (octal PSRAM) |  |
| pass | no signal is assigned to GPIO37 (octal PSRAM) |  |
| pass | no signal is assigned to GPIO45 (VDD_SPI strapping pin) |  |
| pass | the reserved pins are named in the header so nobody re-uses them |  |
| value | GPIOs assigned | 0, 1, 2, 3, 4, 5, 6, 7, 8, 9, 10, 11, 12, 13, 14, 15, 16, 17, 18, 19, 20, 21, 38, 39, 40, 41, 42, 43, 44, 46, 47, 48 |
| note | firmware status | never compiled against a real ESP-IDF installation and never run on hardware; five drivers are stubs (DSN-EEG-003 Rev D section 5) |

### Station 09 functional test TST-EEG-004

|  | Check | Detail |
| --- | --- | --- |
| pass | test specification present |  |
| pass | test fixture design present |  |
| value | layer count | 4 – L1 signal, L2 and L3 reference planes, L4 signal |
| value | input RC corner | 234 Hz |
| pass | E-10: response at 100 Hz is inside +/-1.0 dB | -0.728 dB with the 68 k fitted, against the +/-1.0 dB branch E-10 states for it |
| value | Johnson noise of the 68 k series resistor | 0.280 uV RMS (0.5-70 Hz) |
| pass | E-03: total input-referred noise is inside 1.0 uV RMS | 0.313 uV RMS |
| value | envelope filter | f0 = 49.9 Hz, Q = 0.74, -3 dB at 52.2 Hz |
| value | envelope filter over the C0G tolerance | f0 = 47.5 to 52.5 Hz (-4.8 % to +5.3 %) with C21/C41/C61 and C22/C42/C62 at +/-5 %; Q unchanged, it is set by C2/C1 alone |
| pass | T12e: the calculated envelope corner is inside TST-EEG-004’s 42 to 58 Hz | f0 = 49.9 Hz nominal, measured and recorded per unit |
| pass | E-11: the envelope corner holds 50 Hz +/-10 % over the capacitor tolerance | f0 spans 47.5 to 52.5 Hz against a 45.0 to 55.0 band |
| value | envelope AC-coupling corner | 1.59 Hz |
| pass | E-11 as restated by ECO-EEG-027: AC corner at or below 2 Hz | 1.59 Hz with C20/C40/C60 = 10 uF into 10 kOhm |
| value | envelope output divider | x0.0909 |
| pass | E-11: a 1.1 V peak envelope lands inside +/-100 mV | 100.0 mV |
| value | comparator threshold | 52.1 mV |
| pass | E-12: the comparator trips within the envelope’s working range | 52.1 mV |
| value | contact-light current per site | 1.30 mA |
| pass | E-27, current only: GPIO48 can source all eight lights | 10.4 mA total against a 40 mA rating |
| pass | E-27: the bicolour phase driver exists in the firmware | lights_phase() alternates LED_V against the shift register |
| pass | E-27: all three colours are reachable from what ads_init() enables | green, amber and red are computed from g_loff_insens, g_loff_sens, and the converter is configured to feed a swept LOFF_SENSP comparator |
| open | E-27 has never been seen to light | the driver is written and the current budget is met, but no unit exists, so no light has been driven and TST-EEG-004 T11 – which reads the R/G ratio with a colorimeter – has not been run. The alternation also quantises to the FreeRTOS tick, about 250 Hz rather than exactly 240, which meets the ‘above 100 Hz’ E-27 is written against and is what T11 will actually measure |
| value | raw sample payload at 1000 Hz | 50.0 kB/s (16 channels x 3 bytes + 2 aux, 1000 times a second) |
| value | framed stream at 1000 Hz | 50.8 kB/s (1014 bytes of frame, 1015 after COBS, one every 20 ms) |
| pass | the framed rate matches the 50.7 kB/s ruled in RUL-EEG-021 section B | 50.75 kB/s |
| pass | E-20: one-bit SDMMC has headroom at 1000 Hz | 50.8 kB/s needed against about 2000 kB/s available |
| pass | F-12: full-speed USB carries three times the stream | 152 kB/s against about 1200 kB/s usable |
| value | ring buffer depth at 1000 Hz | 126 s of raw samples (6,291,456 bytes at 50.0 kB/s); 124 s if counted over the framed stream |
| pass | F-06 as relaxed by ECO-EEG-025: at least 90 s of ring | 124 s on the pessimistic framed count; the microSD copy covers anything longer |
| note | the 118 s figure is withdrawn | it divided 6 MB decimal by the framed rate; the ring is 6 MiB and it holds raw samples, so the depth is 126 s (124 s framed) – ECO-EEG-025 |
| value | recording endurance at 1000 Hz | 20.0 h at the 150 mA of TST-EEG-004 T3, but only 6.8 h at the 440 mA that ICD-EEG-006 section 2.7 tallies |
| pass | E-22 on the T3 limit: at least four hours of recording | 20.0 h |
| pass | E-22 also holds at the higher of the two disputed currents | 6.8 h at 440 mA, so E-22 is met on either figure |
| open | the two board-current figures cannot both be right | TST-EEG-004 T3 limits J13 to 150 mA while ICD-EEG-006 section 2.7 tallies about 440 mA. E-22 is met either way (20.0 h against 6.8 h), but the charger, the thermal budget and the T3 limit itself all rest on the number, and it is open item 14 of RFQ-EEG-001 Rev E. It is measured, not calculated, at T3 |
| value | patient auxiliary current, normal condition | 0.106 uA (BAV99 leakage dominates) |
| pass | S-02: normal condition is inside 10 uA DC | 0.106 uA |
| value | single-fault current with a shorted clamp | 36.8 uA, limited by the 68 k series resistor |
| pass | S-02: single fault is inside the 50 uA DC limit | 36.8 uA with the 68 kOhm series resistor fitted; the input corner is 234 Hz, which E-10 covers |
| open | SR-01 is closed in the design and not yet signed off | S-02 is now met at 36.8 uA against 50 uA, because ECO-EEG-024 raised R1-R16 to 68 kOhm. Applying the fix the analysis pointed to is not the same as having it approved: the electrical safety reviewer of RISK-EEG-011 section 7 owns SR-01 and that review has not started. |

### Station 10 software integration with the session runner

|  | Check | Detail |
| --- | --- | --- |
| pass | frame format is fully specified in the RFQ |  |
| value | frame at 1000 Hz | 1015 bytes before COBS, one every 20 ms |
| pass | a 20 ms frame fits comfortably in a full-speed bulk transfer stream | 1015 bytes |
| note | software contract: F-01 composite CDC-ACM + vendor bulk | FW-EEG-001 section 4 |
| note | software contract: F-02 WebUSB and MS OS 2.0 descriptors | FW-EEG-001 section 4 |
| note | software contract: F-04 iSerialNumber = the unit serial TIOV-B-nnnn | PKG-EEG-015 section 5 defines the format; RUL-EEG-021 section B rules it |
| note | software contract: F-08 signed block chain | FW-EEG-001 section 5.6 |
| note | software contract: F-21 timing self-test | TST-EEG-004 step T13 |
| pass | host verification tool covers the signature chain (T16) |  |

### Station 11 final assembly, labelling and kit packing

|  | Check | Detail |
| --- | --- | --- |
| pass | packing, labelling and shipping specification present |  |
| pass | participant quick-start card present |  |
| pass | every foam layer the packing list needs is supplied | CASE-00 Rev C, seven 25 mm layers on a 516 x 390 mm sheet |

### Station 12 despatch

|  | Check | Detail |
| --- | --- | --- |
| pass | regulatory and compliance file present |  |
| pass | lithium shipping is covered (UN3481 / PI967) |  |
| pass | RoHS and REACH declarations are covered |  |

### Station 13 the participant’s own handling

|  | Check | Detail |
| --- | --- | --- |
| pass | quick-start card covers: the one counter-intuitive instruction |  |
| pass | quick-start card covers: what a red light means |  |
| pass | quick-start card covers: charging and the wear-while-charging prohibition |  |
| pass | quick-start card covers: the strap is removable and why |  |
| pass | quick-start card covers: the honesty statement |  |
| note | manual handling | fitting is three steps and is expected to take two to three minutes ONCE THE PARTICIPANT HAS DONE IT ONCE (IFU-EEG-014 section 2). A first-timer is allowed up to five minutes before IFU-EEG-014 is treated as wrong (IFU-EEG-014 fit-trial verification; DSN-EEG-002 section 9) |

### Station 14 return, refurbishment and the next cycle

|  | Check | Detail |
| --- | --- | --- |
| pass | service and refurbishment manual present |  |
| pass | risk analysis covers cross-infection between participants |  |
| pass | change control and document register present |  |