## Kit BOM
Item | Group | Description | Reference part / module | Approved alternate | Qty per kit | Qty @2 | Qty @10 | Qty @25 | Qty @50 | Substitutable | Notes | Bidder unit price @2 | @10 | @25 | @50 | Bidder notes
1 | Modules | ADS1299 8-channel breakout with SPI header and on-board regulators; DAISY_IN and CLKOUT accessible | PIEEG-8 module | generic ADS1299 breakout meeting ICD-EEG-006 section 2.1 | 2 | 4 | 20 | 50 | 100 | No (module class fixed) | Fitted twice per unit. Interface, not brand, is what is specified
2 | Modules | ESP32-S3 development board, 16 MB flash, 8 MB octal PSRAM | Espressif ESP32-S3-DevKitC-1-N16R8 | none | 1 | 2 | 10 | 25 | 50 | No | Inserted directly into J6/J7 on 22.86 mm row spacing. GPIO35, 36, 37 and 45 are not connected on the carrier
3 | Modules | Audio codec module with headphone amplifier, I2S | ES8388 breakout | WM8960 module (firmware change) | 1 | 2 | 10 | 25 | 50 | Yes
4 | Modules | USB 2.0 full-speed isolator module, >= 2.5 kV RMS | ADuM4160 isolator module | none | 1 | 2 | 10 | 25 | 50 | No | OPEN, and NOT closed by the interim answer. The candidate module presents USB-B and E-24 asks for USB-C; the WH-09 panel pigtail is the interim answer and E-24 says of it, in terms, 'it is not a settled design'. WHAT IS MISSING is an ADuM4160-class isolator module with a USB-C host connector, qualified against E-24 and the DSN-EEG-003 section 3.3 isolation keep-out. Until one exists the non-conformance is live and is carried, not resolved. Bidders who know of such a module should say so
5 | Modules | Secure element breakout, I2C | Adafruit ATECC608 breakout (4314) | any ATECC608B breakout | 1 | 2 | 10 | 25 | 50 | No
6 | Modules | Li-ion charger with power path and CE pin | Adafruit bq24074 (4755) | MCP73871 module | 1 | 2 | 10 | 25 | 50 | Yes | Combined charger-plus-gauge assembly is the baseline
7 | Modules | Battery fuel gauge, I2C | Adafruit MAX17048 (5580) | BQ27441 module | 1 | 2 | 10 | 25 | 50 | Yes | May be combined with the charger on one board
8 | Modules | Buck-boost regulator module, VSYS to 5.00 V at 1 A | TPS63020 breakout | TPS63070 module | 1 | 2 | 10 | 25 | 50 | Yes | NEW IN v2. Package v1 had no source for the 3.3 V rail at all (ECO-EEG-002)
9 | Modules | microSD breakout, SDMMC | generic microSD SDMMC breakout | none | 1 | 2 | 10 | 25 | 50 | Yes | One-bit SDMMC. 70 kB/s needed, about 2 MB/s available
10 | Modules | Electret microphone preamplifier module, fixed gain | MAX4466 class (Adafruit 1063) | none -- specified by interface, ICD-EEG-006 section 2.9 | 1 | 2 | 10 | 25 | 50 | Yes | OPEN, but narrower than it was. The AGC CONFLICT IS RESOLVED: the MAX9814 of package v1 is an AGC part, is NOT APPROVED, and is superseded here -- E-14 requires AGC off, and the fixed-gain MAX4466 class named here is E-14's own preferred route. What is still open is the SELECTION, not the conflict -- E-14 reads 'which preamplifier is not settled' and specifies the module BY INTERFACE in ICD-EEG-006 section 2.9 until one is bought and measured. Quote against the interface. Mounts on MP-01 at J21; the boom carries the bare capsule
11 | Modules | Room microphone module with hardware mute | to be selected | none | 1 | 2 | 10 | 25 | 50 | Yes | OPEN, and the only row in this BOM with NO PART AT ALL. E-15 requires the audio path to be hardware-gatable on MIC_MUTE (GPIO21); a catalogue search found no breakout with a mute pin -- the common electret and MEMS boards (SparkFun, Pololu, Adafruit 1063, INMP441) bring out audio and power only. The fallback is a capsule plus a TS5A3159 analogue switch on a programme-designed adapter, which is a NEW DRAWING and has not been drawn. Thirteen module assemblies per unit and this is the unspecified one. Bidders: quote the other twelve and say whether you would build the adapter to a drawing when one exists
12 | Modules | 8-bit shift register module | 74HC595 breakout | none | 1 | 2 | 10 | 25 | 50 | Yes | Q0..Q7 must be brought out. In v1 they were not, and the contact lights had no driver at all (ECO-EEG-001)
13 | Modules | Protected 18650 Li-ion cell >= 3000 mAh, UN 38.3 | Panasonic NCR18650B protected | Samsung 35E in a protection holder | 1 | 2 | 10 | 25 | 50 | Yes | UN 38.3 test summary required with the first delivery
14 | Modules | microSD card 32 GB high endurance | SanDisk Industrial | Kingston SDCIT2 | 1 | 2 | 10 | 25 | 50 | Yes
15 | Carrier | EEG-CAR-01 Rev B carrier PCB, FOUR layers, 150.0 x 130.0 mm, ENIG | fabricated from the released Gerber set | none | 1 | 2 | 10 | 25 | 50 | n/a | v1 specified two layers and 130 x 124 mm; see DSN-EEG-003 Rev D section 3
16 | Carrier | EEG-CAR-01 assembly per the carrier sub-BOM sheet | see sheet 'Carrier sub-BOM' | per sub-BOM alternates | 1 | 2 | 10 | 25 | 50 | Per sub-BOM
17 | Carrier | Ribbon-jumper set, keyed, per ICD-EEG-006 section 3 | 2.54 mm IDC ribbon, made up | none | 1 | 2 | 10 | 25 | 50 | Yes | Maximum 60 mm on the analogue jumper
18 | Printed | HM-01 helmet frame monocoque, PA12 MJF | MJF PA12, dyed graphite | SLS PA12 | 1 | 2 | 10 | 25 | 50 | Process only | About 240 g
19 | Printed | HM-04 electrode assembly body | MJF PA12 | SLS PA12 | 10 | 20 | 100 | 250 | 500 | Process only | Eight fitted, two spare
20 | Printed | HM-08 battery hatch (called HM-07 in v1 -- see PARTS-EEG-019) | MJF PA12 | SLS PA12 | 1 | 2 | 10 | 25 | 50 | Process only
21 | Printed | HM-07 boom microphone arm | MJF PA12 + gooseneck | none | 1 | 2 | 10 | 25 | 50 | Process only
22 | Printed | MP-01 module mounting plate | MJF PA12 | none | 1 | 2 | 10 | 25 | 50 | Process only | NEW IN v2. 146.0 x 126.0 x 3.0 mm
23 | Printed | POD-P1 enclosure base and lid -- Phase 1 only | MJF PA12 or FDM PETG | none | 1 | 2 | 10 | 25 | 50 | Process only | 163.0 x 143.0 x 58.0 mm external
24 | Printed | HM-02 TPU comfort pads x4 and chin-cup liner | TPU 85A | cast | 1 | 2 | 10 | 25 | 50 | Yes | Consumable, replaced every turnaround
25 | Printed | HM-09 service key | MJF PA12 | none | 0 | 0 | 0 | 0 | 0 | Process only | ONE PER OPERATOR, deliberately NOT in the participant's kit
26 | Printed | FIT-01 fit-test coupon | same build as the batch | none | 0 | 0 | 0 | 0 | 0 | Process only | One per print batch, checked before the batch is accepted
27 | Bought-in mechanical | Electrode assembly springs 3-6 N, stainless | per HM-04 drawing | none | 10 | 20 | 100 | 250 | 500 | Yes
28 | Bought-in mechanical | Occipital yoke ratchet assembly (POM pawl) | hard-hat style ratchet | none | 1 | 2 | 10 | 25 | 50 | Yes | 2 mm per click, 52-62 cm
29 | Bought-in mechanical | Boom gooseneck 120 mm and mount | generic | none | 1 | 2 | 10 | 25 | 50 | Yes
30 | Bought-in mechanical | Chin strap webbing, chin cup and anchors | generic | none | 1 | 2 | 10 | 25 | 50 | Yes | Removable. The runner says so before it is first mentioned
31 | Bought-in mechanical | M3 x 18 nylon hex standoffs and M3 x 6 screws | Wurth 970200321 class | none | 4 | 8 | 40 | 100 | 200 | Yes | Carrier to MP-01
32 | Bought-in mechanical | Helmet harness WH-01, 12-way screened, per WH-EEG-008 | custom assembly | none | 1 | 2 | 10 | 25 | 50 | Yes | v1 ran the light lines through this cable; v2 splits them out (ECO-EEG-014)
33 | Bought-in mechanical | Contact-light harness WH-02, 10-way, per WH-EEG-008 | custom assembly | none | 1 | 2 | 10 | 25 | 50 | Yes | NEW IN v2
34 | Electrodes | Sintered Ag/AgCl cup electrodes on service bayonet | Wuhan Greentek sintered cup | Florida Research Instruments / Spes Medica | 10 | 20 | 100 | 250 | 500 | Yes | Replaced outright every ~25 sessions
35 | Electrodes | Ag/AgCl ear-clip reference electrodes | Greentek ear clip | none | 2 | 4 | 20 | 50 | 100 | Yes
36 | Electrodes | EMG snap leads to DIN 42802 plug, 1 m | Greentek / generic | none | 3 | 6 | 30 | 75 | 150 | Yes
37 | Electrodes | Disposable EMG snap electrodes, pack of 30 | Ambu BlueSensor N | Kendall | 1 | 2 | 10 | 25 | 50 | Yes
38 | Electrodes | Bicolour contact-light LEDs, two-lead red/green | generic 3 mm | none | 10 | 20 | 100 | 250 | 500 | Yes | NEW IN v2. Two-lead inverse-parallel, driven in two phases at 240 Hz
39 | Audio | Closed-back over-ear headphones, 32 to 64 ohm, 3.5 mm | Audio-Technica ATH-M20x (47 ohm) | Superlux HD-681 | 1 | 2 | 10 | 25 | 50 | Yes | A-04 restated to 32-64 ohm; the calibrated level is measured per model
40 | Audio | Electret capsule and windscreen for the boom | Primo EM272 | generic | 1 | 2 | 10 | 25 | 50 | Yes | The capsule only; its preamplifier is on MP-01
41 | Consumables | EEG paste 100 g x2, abrasive prep gel 100 g, saline wipes x30, blunt-tip syringes x4 | Greentek GT20/GT5 | Ten20 / NuPrep | 1 | 2 | 10 | 25 | 50 | Yes
42 | Cables | USB-C to USB-A 1 m and USB-C to USB-C 1 m | certified generic | none | 1 | 2 | 10 | 25 | 50 | Yes | One of them is the host lead. There is no captive cable in Phase 1
43 | Cables | 5 V 2 A USB charger, EU plug, CE | certified generic | none | 1 | 2 | 10 | 25 | 50 | Yes | The helmet is never worn while the charge cable is connected
44 | Case | Hard-shell IP67 case, internal >= 516 x 390 mm and >= 210 mm deep, of which >= 185 mm in the base | Peli 1560, unwheeled variant (published internal 518 x 392 x 229 mm) | none | 1 | 2 | 10 | 25 | 50 | Yes | Sized for an assembled HM-01 standing upright on the halo, beside the POD-P1 bay. PKG-EEG-015 section 3.2 governs the envelope; Nanuk 915 and Peli 1450 are withdrawn as about 155 mm deep. Larger than RFQ M-05, see CN-PKG-03. OPEN: no shell has been bought or measured; the figures are the manufacturer's published ones
45 | Case | Seven-layer PE foam insert, 516 x 390 mm sheet, 25 mm per layer, 175 mm stack, nine bays | die-cut or laser-cut to CASE-00 Rev C | pluck foam (Phase 1) | 1 | 2 | 10 | 25 | 50 | Yes | OPEN: the CASE-00 Rev C cut file has not been drawn (PKG-EEG-015 section 2.2 and its open item 1, CN-PKG-02). The two Rev B DXF layers supplied with package v1 are superseded and must not be cut. Phase 1 packs in pluck foam
46 | Case | Laminated quick-start card and packing photograph; double-wall outer carton | print | none | 1 | 2 | 10 | 25 | 50 | Yes | IFU-EEG-014 is the artwork
47 | Service | Provisioning, firmware load and TST-EEG-004 Rev C | per TST-EEG-004 | none | 1 | 2 | 10 | 25 | 50 | n/a | About 25 minutes per unit, 20 of them unattended
Rev C, 1 September 2026. Four-layer carrier, 150.0 x 130.0 mm, in the enlarged POD-P1. Supersedes Rev B entirely. Yellow cells are for the bidder; no prices are given in this file by design. Nothing in this package has been manufactured or measured.

## Carrier sub-BOM (generated)
Item | Qty | Designators | Value | Footprint | Manufacturer part number | Fit | Notes | Bidder unit price
1 | 19 | C1 C2 C3 C4 C5 C6 C7 C8 C9 C10 C11 C12 C13 C14 C15 C16 C21 C41 C61 | 10n C0G 50V | C_0603_1608Metric | Murata GCM1885C1H103JA16D | fit | RF filter to AGND_REF, Fz scalp
2 | 12 | C20 C40 C60 C70 C71 C72 C73 C74 C89 C90 C100 C101 | 10u X5R 16V | C_0603_1608Metric | Murata GRM188R61C106MA73D | fit | input AC coupling into 10 kOhm: 1.6 Hz corner, envelope channel 1. ECO-EEG-027: 1 uF gave 15.9 Hz, which removes the envelope it is meant to pass.
3 | 3 | C22 C42 C62 | 22n C0G 50V | C_0603_1608Metric | Murata GCM1885C1H223JA16D | fit | Sallen-Key feedback C, sets Q = 0.74, channel 1. C0G for the same reason as C21; the ratio 22/10 is what holds Q at 0.742.
4 | 18 | C23 C24 C43 C44 C50 C51 C52 C63 C64 C80 C81 C82 C83 C84 C85 C86 C87 C88 | 100n X7R 25V | C_0603_1608Metric | Murata GCM188R71E104KA57D | fit | AVDD decoupling at U1
5 | 2 | C102 C103 | 100n X7R 50V | C_0603_1608Metric | Murata GCM188R71H104KA57D | fit | AVDD high-frequency decoupling
6 | 17 | D1 D2 D3 D4 D5 D6 D7 D8 D9 D10 D11 D12 D13 D14 D15 D16 D23 | BAV99 | SOT-23 | Nexperia BAV99,215 | fit | clamp to AVDD/AVSS, Fz scalp
7 | 3 | D20 D40 D60 | BAT54S | SOT-23 | Nexperia BAT54S,215 | fit | precision-rectifier Schottky pair (pin 3 at the op-amp output), channel 1
8 | 1 | D24 | PESD5V0S2BT | SOT-23 | Nexperia PESD5V0S2BT,215 | fit | transient suppressor on the charge input
9 | 3 | FID1 FID2 FID3 | global fiducial | Fiducial_1mm_Mask3mm | - | fit | 1 mm copper, 3 mm mask opening, no net
10 | 1 | F1 | PTC 1.1 A hold / 2.2 A trip | R_1206_3216Metric | Bourns MF-MSMF110-2 | fit | resettable fuse in the charge input
11 | 1 | J1 | ADS1299 module #1 digital | PinSocket_1x12_P2.54mm_Vertical | Samtec SSW-112-01-G-S or equivalent 2.54 mm socket strip | fit | 1x12 socket: SPI, control lines, shared clock, 3V3 and 5 V feed to module #1
12 | 1 | J2 | ADS1299 module #1 analogue signals | PinSocket_1x10_P2.54mm_Vertical | Samtec SSW-110-01-G-S or equivalent 2.54 mm socket strip | fit | 1x10 socket: eight EEG inputs, the SRB1 reference and the bias drive
13 | 1 | J3 | ADS1299 module #2 digital | PinSocket_1x12_P2.54mm_Vertical | Samtec SSW-112-01-G-S or equivalent 2.54 mm socket strip | fit | 1x12 socket: SPI, control lines, shared clock, 3V3 and 5 V feed to module #2
14 | 1 | J4 | ADS1299 module #2 analogue signals | PinSocket_1x10_P2.54mm_Vertical | Samtec SSW-110-01-G-S or equivalent 2.54 mm socket strip | fit | 1x10 socket: three EMG inputs, three envelope channels, two spares, SRB1
15 | 1 | J5 | ADS #1 DAISY_IN / CLKOUT stub | PinSocket_1x04_P2.54mm_Vertical | Samtec SSW-104-01-G-S or equivalent 2.54 mm socket strip | fit | 1x4 stub: module #2 DOUT into module #1 DAISY_IN and the shared 2.048 MHz clock
16 | 1 | J6 | ESP32-S3-DevKitC-1 row A | PinSocket_1x22_P2.54mm_Vertical | Samtec SSW-122-01-G-S or equivalent 2.54 mm socket strip | fit | 1x22 socket, DevKitC-1 left header. Row spacing to J7 is 22.86 mm (0.900 in).
17 | 1 | J7 | ESP32-S3-DevKitC-1 row B | PinSocket_1x22_P2.54mm_Vertical | Samtec SSW-122-01-G-S or equivalent 2.54 mm socket strip | fit | 1x22 socket, DevKitC-1 right header
18 | 1 | J8 | ES8388 codec module | PinSocket_1x14_P2.54mm_Vertical | Samtec SSW-114-01-G-S or equivalent 2.54 mm socket strip | fit | 1x14: I2S, I2C, headphone amplifier output, the stimulus envelope tap and 5 V
19 | 1 | J9 | Codec microphone feeds | PinSocket_1x04_P2.54mm_Vertical | Samtec SSW-104-01-G-S or equivalent 2.54 mm socket strip | fit | 1x4: voice and room preamp outputs into the codec ADC, and the room-mic mute line
20 | 1 | J10 | ADuM4160 USB isolator, device side | PinSocket_1x04_P2.54mm_Vertical | Samtec SSW-104-01-G-S or equivalent 2.54 mm socket strip | fit | 1x4: isolated USB D+/D-, isolated 3V3 and DGND. No carrier copper crosses the barrier. The module's host receptacle is USB-B on the qualified part, not USB-C; WH-09 adapts it. See RUL-EEG-021 section B.
21 | 1 | J11 | ATECC608B breakout | PinSocket_1x04_P2.54mm_Vertical | Samtec SSW-104-01-G-S or equivalent 2.54 mm socket strip | fit | 1x4: secure element on the shared I2C bus
22 | 1 | J12 | Charger + fuel gauge module | PinSocket_1x08_P2.54mm_Vertical | Samtec SSW-108-01-G-S or equivalent 2.54 mm socket strip | fit | 1x8: bq24074-class charger with power path, and the MAX17048 gauge on I2C
23 | 1 | J13 | 18650 protected cell | JST_PH_B2B-PH-K_1x02_P2.00mm_Vertical | JST B2B-PH-K-S(LF)(SN) | fit | battery input from the keyed cell carrier
24 | 1 | J14 | Helmet harness -- electrodes (12-way screened) | PinSocket_1x12_P2.54mm_Vertical | Samtec SSW-112-01-G-S or equivalent 2.54 mm socket strip | fit | 1x12 socket: eight scalp electrodes, two ear references, the bias lead and the cable screen. Nothing digital enters this connector.
25 | 1 | J15 | DIN 42802 EMG1 cheek | DIN42802_1p5mm_Socket | Staubli SLB1,5-F / LB-I1,5 | fit | touch-proof 1.5 mm safety socket, EMG1 cheek, colour-coded
26 | 1 | J16 | DIN 42802 EMG2 submental | DIN42802_1p5mm_Socket | Staubli SLB1,5-F / LB-I1,5 | fit | touch-proof 1.5 mm safety socket, EMG2 submental, colour-coded
27 | 1 | J17 | DIN 42802 EMG3 laryngeal | DIN42802_1p5mm_Socket | Staubli SLB1,5-F / LB-I1,5 | fit | touch-proof 1.5 mm safety socket, EMG3 laryngeal, colour-coded
28 | 1 | J18 | Boom microphone pigtail (TRRS) | PinSocket_1x04_P2.54mm_Vertical | Samtec SSW-104-01-G-S or equivalent 2.54 mm socket strip | fit | 1x4: to the panel TRRS jack at the left temple
29 | 1 | J19 | 74HC595 contact-light driver | PinSocket_1x16_P2.54mm_Vertical | Samtec SSW-116-01-G-S or equivalent 2.54 mm socket strip | fit | 1x16: shift-register module. Q0..Q7 drive the eight helmet contact lights.
30 | 1 | J20 | microSD breakout, SDMMC 1-bit | PinSocket_1x08_P2.54mm_Vertical | Samtec SSW-108-01-G-S or equivalent 2.54 mm socket strip | fit | 1x8: microSD card breakout wired for one-bit SDMMC (70 kB/s needed, 2 MB/s available)
31 | 1 | J21 | Boom-mic preamp module (part not settled, E-14) | PinSocket_1x06_P2.54mm_Vertical | Samtec SSW-106-01-G-S or equivalent 2.54 mm socket strip | fit | 1x6: fixed-gain boom microphone preamp module. RFQ E-14 forbids automatic gain control, so the MAX9814 of package v1 is NOT APPROVED (AVL-EEG-017); a MAX4466-class fixed-gain board is the reference. The part is an open item and this socket is specified by its interface, not by a chosen module.
32 | 1 | J22 | EOG / spare electrode header | PinSocket_1x03_P2.54mm_Vertical | Samtec SSW-103-01-G-S or equivalent 2.54 mm socket strip | fit | 1x3 socket: two spare protected electrode channels and their screen. Wired to panel DIN sockets only when the EOG option is used (RFQ 3.1).
33 | 1 | J23 | ADS1299 module #1 analogue rails | PinSocket_1x06_P2.54mm_Vertical | Samtec SSW-106-01-G-S or equivalent 2.54 mm socket strip | fit | 1x6 socket: AVDD, AVSS, BIASIN, AGND_REF and a second rail pair
34 | 1 | J24 | Charge input from the panel USB-C | JST_PH_B2B-PH-K_1x02_P2.00mm_Vertical | JST B2B-PH-K-S(LF)(SN) | fit | charge-only USB-C receptacle pigtail; no data conductor enters here
35 | 1 | J25 | Buck-boost module, TPS63020 class | PinSocket_1x06_P2.54mm_Vertical | Samtec SSW-106-01-G-S or equivalent 2.54 mm socket strip | fit | 1x6: VSYS to a regulated 5.0 V rail for the DevKit and both ADS1299 modules
36 | 1 | J26 | Debug / programming header | PinSocket_1x06_P2.54mm_Vertical | Samtec SSW-106-01-G-S or equivalent 2.54 mm socket strip | fit | 1x6: 3V3, GND, UART0 TX and RX, EN and NC_GPIO0 -- console and recovery flashing. Way 6 is the spare NC_GPIO0, not GPIO0: GPIO0 is committed to LED_SR_LATCH.
37 | 1 | J27 | Headphone jack pigtail | PinSocket_1x04_P2.54mm_Vertical | Samtec SSW-104-01-G-S or equivalent 2.54 mm socket strip | fit | 1x4: to the panel 3.5 mm headphone jack
38 | 1 | J28 | Room microphone module | PinSocket_1x04_P2.54mm_Vertical | Samtec SSW-104-01-G-S or equivalent 2.54 mm socket strip | fit | 1x4: outward-facing room microphone module with hardware mute
39 | 1 | J29 | ADS1299 module #2 analogue rails | PinSocket_1x06_P2.54mm_Vertical | Samtec SSW-106-01-G-S or equivalent 2.54 mm socket strip | fit | 1x6 socket: AVDD2, AVSS2, BIASIN, AGND_REF and a second rail pair
40 | 1 | J30 | Helmet harness -- contact lights (10-way) | PinSocket_1x10_P2.54mm_Vertical | Samtec SSW-110-01-G-S or equivalent 2.54 mm socket strip | fit | 1x10 socket: eight contact-light lines, the light common and its return. This is the second helmet cable; it carries no electrode signal.
41 | 1 | L1 | 600R at 100 MHz, 500 mA | L_0603_1608Metric | Murata BLM18AG601SN1D | fit | ferrite between DVDD3V3 and the isolator device-side supply
42 | 16 | R1 R2 R3 R4 R5 R6 R7 R8 R9 R10 R11 R12 R13 R14 R15 R16 | 68k 0.1% 25ppm | R_0603_1608Metric | Vishay TNPW060368K0BEEA | fit | series protection resistor, Fz scalp
43 | 12 | R20 R21 R23 R24 R40 R41 R43 R44 R60 R61 R63 R64 | 10k 0.1% | R_0603_1608Metric | Vishay TNPW060310K0BEEA | fit | rectifier input resistor, envelope channel 1
44 | 3 | R22 R42 R62 | 4k99 0.1% | R_0603_1608Metric | Vishay TNPW06034K99BEEA | fit | half-wave summing resistor (R/2) into the absolute-value stage, channel 1
45 | 6 | R25 R26 R45 R46 R65 R66 | 215k 0.1% | R_0603_1608Metric | Vishay TNPW0603215KBEEA | fit | Sallen-Key R1, channel 1
46 | 3 | R27 R47 R67 | 22k 0.1% | R_0603_1608Metric | Vishay TNPW060322K0BEEA | fit | output divider, top, channel 1
47 | 3 | R28 R48 R68 | 2k2 0.1% | R_0603_1608Metric | Vishay TNPW06032K20BEEA | fit | output divider, bottom -- scales to +/-100 mV, channel 1
48 | 6 | R50 R51 R52 R81 R83 R88 | 10k 1% | R_0603_1608Metric | Vishay CRCW060310K0FKEA | fit | pull-up for BTN_A
49 | 8 | R70 R71 R72 R73 R74 R75 R76 R77 | 1k 1% | R_0603_1608Metric | Vishay CRCW06031K00FKEA | fit | contact-light series resistor, helmet site 1
50 | 7 | R78 R79 R87 R90 R91 R92 R93 | 0R | R_0603_1608Metric | Vishay CRCW06030000Z0EA | fit | LED_V common drive link from GPIO48. 0R fitted; 47R is the approved alternate.
51 | 1 | R80 | 470k 1% | R_0603_1608Metric | Vishay CRCW0603470KFKEA | fit | comparator threshold divider, top (trips at about 52 mV)
52 | 1 | R82 | 1M 1% | R_0603_1608Metric | Vishay CRCW06031M00FKEA | fit | comparator hysteresis, about 5 mV
53 | 2 | R84 R86 | 100k 1% | R_0603_1608Metric | Vishay CRCW0603100KFKEA | fit | VBUS-present divider, top
54 | 1 | R85 | 150k 1% | R_0603_1608Metric | Vishay CRCW0603150KFKEA | fit | VBUS-present divider, bottom. ECO-EEG-022: 3.00 V at VBUS = 5 V, above the 2.48 V input-high threshold. The 56 kOhm of the first cut gave 1.79 V and would not have asserted reliably.
55 | 1 | R89 | 2k2 1% -- DNP | R_0603_1608Metric | Vishay CRCW06032K20FKEA | DNP | electret bias for the boom capsule. DO NOT POPULATE unless the preamp module lacks its own microphone bias; see ICD-EEG-006 section 7.2.
56 | 2 | R94 R95 | 4k7 1% | R_0603_1608Metric | Vishay CRCW06034K70FKEA | fit | I2C SDA pull-up to DVDD3V3
57 | 1 | SW1 | BTN_A response (green) | SW_PUSH_6mm_H7.3mm | Omron B3F-1052 | fit | response button A, 12 mm green cap on an extender
58 | 1 | SW2 | BTN_B response (blue) | SW_PUSH_6mm_H7.3mm | Omron B3F-1052 | fit | response button B, 12 mm blue cap on an extender
59 | 1 | SW3 | BTN_STOP (red) | SW_PUSH_6mm_H7.3mm | Omron B3F-1052 | fit | session stop button, 12 mm red cap, distinct tactile feel
60 | 1 | TP1 | SCLK | TestPoint_Pad_D1.5mm | - | fit | test point, SCLK
61 | 1 | TP2 | MOSI | TestPoint_Pad_D1.5mm | - | fit | test point, MOSI
62 | 1 | TP3 | MISO | TestPoint_Pad_D1.5mm | - | fit | test point, MISO
63 | 1 | TP4 | DRDY | TestPoint_Pad_D1.5mm | - | fit | test point, DRDY
64 | 1 | TP5 | I2S_BCLK | TestPoint_Pad_D1.5mm | - | fit | test point, I2S_BCLK
65 | 1 | TP6 | I2S_LRCK | TestPoint_Pad_D1.5mm | - | fit | test point, I2S_LRCK
66 | 1 | TP7 | ENV_STIM | TestPoint_Pad_D1.5mm | - | fit | test point, stimulus envelope
67 | 1 | TP8 | ENV_VOICE | TestPoint_Pad_D1.5mm | - | fit | test point, voice envelope
68 | 1 | TP9 | ENV_ROOM | TestPoint_Pad_D1.5mm | - | fit | test point, room envelope
69 | 1 | TP10 | AVDD | TestPoint_Pad_D1.5mm | - | fit | test point AVDD
70 | 1 | TP11 | AVSS | TestPoint_Pad_D1.5mm | - | fit | test point AVSS
71 | 1 | TP12 | DVDD3V3 | TestPoint_Pad_D1.5mm | - | fit | test point, DVDD3V3
72 | 1 | TP13 | AGND_REF | TestPoint_Pad_D1.5mm | - | fit | test point AGND_REF
73 | 1 | TP14 | DGND | TestPoint_Pad_D1.5mm | - | fit | test point, DGND
74 | 1 | TP15 | VSYS | TestPoint_Pad_D1.5mm | - | fit | test point, VSYS
75 | 1 | TP16 | V5V | TestPoint_Pad_D1.5mm | - | fit | test point, V5V
76 | 1 | TP17 | NC_DRDY2 | TestPoint_Pad_D1.5mm | - | fit | test point, NC_DRDY2
77 | 1 | TP18 | CLK_ADS | TestPoint_Pad_D1.5mm | - | fit | test point, CLK_ADS
78 | 3 | U1 U2 U3 | OPA4376AIPW | TSSOP-14_4.4x5mm_P0.65mm | TI OPA4376AIPWR | fit | quad precision op-amp: rectifier, absolute-value summer, 50 Hz filter and output buffer for envelope channel 1 (stimulus, from the headphone tap)
79 | 1 | U7 | TLV3201AIDBV | SOT-23-5 | TI TLV3201AIDBVR | fit | stimulus-envelope comparator (RFQ E-12); its output is latched on DRDY in firmware
GENERATED from package_v2.4/tools/design.py by tools/emit_workbooks.py. Do not edit by hand: edit design.py and regenerate, so this sheet and the board can never disagree.

## Read me
EEG field kit -- kit bill of materials
Revision C, 1 September 2026. Companion to RFQ-EEG-001 Rev E and DSN-EEG-003 Rev D.
What changed from Rev B
The carrier is four layers and 150.0 x 130.0 mm, not two layers and 130 x 124 mm.
A buck-boost module is added: package v1 had no source for the 3.3 V rail at all.
The contact lights are two-lead bicolour LEDs and the 74HC595 outputs now reach the carrier; in v1 the lights had no driver.
The helmet harness is two cables, not one.
MP-01, the module mounting plate, is new, and so are the ribbon jumpers.
HM-07 is the boom arm and HM-08 is the battery hatch; v1 used HM-07 for both.
The headphone impedance is restated as 32 to 64 ohm.
The case and the foam insert carry the two corrections PKG-EEG-015 section 1.2 asked for and Rev B did not apply: the internal envelope is >= 516 x 390 mm and >= 210 mm deep, Nanuk 915 and Peli 1450 are withdrawn as too shallow, and the insert is a seven-layer stack, not the two DXF layers v1 supplied.
Five items are OPEN and are marked as such: the isolator module's connector, the boom preamplifier, the room-microphone module, the travel case (no shell has been bought or measured) and the foam insert (the CASE-00 Rev C cut file has not been drawn).
Nothing in this package has been manufactured or measured, and no safety engineer has reviewed the design.